## Lists
Label | Severity | Status | Related to: | Yes/No
closed | High | Resolved | Hardware | Yes
Backlog | Low | Ready for testing | Application | No
CurrentPlan |  | In progress | Documentation | NA
newRequirement |  | No fix required | Yocto
changeRequirement |  | Require inputs from Flexi | Multiple
Not able to reproduce
Unassigned
Failed Testing

## query
ID# | Status | Brief description | Detailed Description | Priority | Items(s) affected | Status2 | Software version with the issue | Who is responsible | Analysis | Intended fix version | Column1 | Needed for Aug Ux Study? | Review Remarks - Aug 4th 2026 | Bug closure dates | S/W App Released date
1 | closed | LCD Flickering | In multiple power cycles the LCD seen flickering. | High | Hardware | Resolved | 0.0.6 | Ayush Pareek;#16 | We have not seen this issue on latest hardware. Since fixing the voltage rail shorting issue. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
3 | closed | Error Low Disk Space | Upon bootup device shows low disk space error | Low | Application | Resolved | 1.1 | davis.packiaraj;#18 |  | 1.3 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
4 | closed | Media Playback not working | Screen Recording works but when we try to play media from gallery the recording does not work | Low | Application | Resolved | 1.3 | davis.packiaraj;#18 | Couldn't replicate the behaviour here. When recording is ON navigating to gallery will be restricted as part of new requirement. So this request is void. Ayush 290725: Found that the device recording does not happen when microphone is enabled. After fixing the pulse audio package in build release 1.3.0, this bug has been fixed | Build 1.3.0 and App Version 0.2.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
5 | closed | Build Version | There is no build version displayed on the Application. It requires to identify the Linux build | Low | Yocto | Resolved | n/a | davis.packiaraj;#18 | Ayush 29-July-2025 : The version text file "image-version" added. That can be used to track the build version. Further this can be looked with new GUI v6 | Build 1.3.0 and Application 0.2.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
6 | closed | Application Crashing | During running the random operation with the device like recording, camera insertion and removal. the application crashed. | Low | Application | Resolved | 1.3 | davis.packiaraj;#18 | Do we have any logs at the time of crash? Crashing happens when device is left un attended after timeout. Bug fixed on version 1.4 |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
7 | closed | Audio Playback | There is a delay or lag between the button press and audio tone playback | Low | Application | Resolved | 1.3 | davis.packiaraj;#18 | Moved to QSoundEffects from QMediaPlayer, which couldn't integrate to pulse audio. Need to test with sdl2 library. PAPLAY worked with low latency and it resolved the lag issue. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
8 | closed | Zoom feature | Currently Zoom feature having 100% , 150% and 200% , SRS shows the requirement for 0.5 ,1 and 2 , does it need to be change or its is going to be same ? <05Sep WW>: SRS/GUI-spec/DIDO to be updated. Software is correct. | Low | Application | Resolved | 1.3 | Abhishek Shukla;#24 | 0.5 is eliminated from scope, instead 1.5 is added. Which is 100%, 150%, 200%. No change required. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
9 | closed | File Operations | Progress Bar for copying the file to external media (USB) not working properly. 1. When copying videos, it takes a long time to copy, but the progress bar does not see to get updated? 2. When copying folders (from a selection in gallery view), the progress bar (blue portion) is not aligned with the grey background portion. <WW 03Sep>: To add more information to this - it took approx. 2min to copy a 49MByte video file from the display to a USB2.0 usb stick, but the same file on the same stick when connected to my laptop took only <5sec to copy to my laptop? <WW 01Jul2026>: I am using a USB2.0 memory stick and it takes time to copy (expected), but the progress bar is stuck on 0%. Copy does complete, but no progress indication during copy. See video on TEAMS 'v0.5.9 issues/Copy to USB progress bar not updating' | Low | Application | Resolved | Yocto_v2.1.9/App_v0.4.0 Yocto_v3.2.9/App_v0.5.9 | davis.packiaraj;#18 | Need to change ADB to USB for testing. Davis: Timer based progress update implemented. Ayush17Sep2025: 1. I can see the progress bar fixed 2. The time taken to copy 232 MB is around 9 minute which is very high. For same data copy with windows laptop to same USB, time take is 8 seconds Davis 24Sep2025: Copied 853.5 MB file which takes 1 Min 15 sec to complete transfer including sync (Transfer completes in 55 sec). From terminal it takes around 56 sec for the same file. Ayush250925: Poornesh will test with the serial terminal Ayush260925: Poornesh found fix, Ayush waiting for release to be done to test this Ayush290925: Tested today and found still it not fixed. Results shared with poornesh Ayush300925: Transfer speed improved but When I try transfer larger file which is 628 MB, the progress bar shows negative number | App: 0.4.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
10 | closed | File Operations | Multiple file copy and delete functions not working | High | Application | Resolved | 1.3 | davis.packiaraj;#18 | Multiple selection bug fixed for delete and copy. | 0.2.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
11 | closed | Snapshot | Snapshot capturing the flashing boarder as well and also capturing the blur image if camera is moving | Medium | Application | Resolved | 1.3 | davis.packiaraj;#18 | Blur is not under application software control. Is it ok to add delay to capture screen? ex: flash border and after flashing is complete, capture the screen. - pending Ayush: Closing this as it is repeat of other | 0.2.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
12 | closed | Account creation failed on first bootup | Software Failed to create the admin account after first bootup. ( File system creation failed ) | Low | Yocto | Resolved | 1.6 | poornesh.g;#21 |  |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
13 | closed | Application Crashing (if increase black level to max) | In application version 0.1.8, tried to adjust the black level to max and the application got a reset | High | Application | Resolved | 0.1.8 | davis.packiaraj;#18 |  | 0.1.9 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
14 | closed | Application crash with sleep button function | It is observed that application got a crash as it stops feeding the fpga watchdog and recover when pressing the power button | High | Application | Resolved |  | Ayush Pareek;#16 | In version RC-09, the issue is fixed. Need to test with new process and the build updated. Current Observation, the brightness get reduced when we press the sleep button when camera is connected | 0.1.8 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
15 | closed | Date selection 31 days in feb month | On the date selection, the feb month is having selection of 31 days. Also, be careful about the condition where if the month was set at e.g. March and date at 31, then the month is decremented back to February, the date must also be changed to limit for the month. | Medium | Application | Resolved | 1.0.8 | davis.packiaraj;#18 | Ayush:022Sept2025: If we select the month as march and set date to 31 and then change month back to feb and then save. The device does not set the date and time but no error message shown in this case. Also, if we switch to feb, after selecting the day as 31 then day should be switch back to max day of the selected month. Davis: Fixed in v0.3.6 <WW 01Jul2026> : Confirmed fixed in version "Yocto_v3.2.9/App_v0.5.9" | App:0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
16 | closed | Software Crashing ¦ Changing LCD Brightness Level to 100 | When we change brightness to 100 then after few seconds device get blank. Recovers if we do powercycle but the it will get blank again until we reduce the LCD Brightness | High | Hardware | Resolved | 1.0.8 | poornesh.g;#21 | This is a hardware issue. Not related to software. Not observed on the hardware i am testing. Need to discuss. Ayush 29-07-25: It is found that, if we connect battery no reset occur. There are resistors identified in hardware to fix it. Analysis is on going. WW:27Aug2025: This is now fixed - BQ25672 charger chip need to be correctly configured as it had current limit set too low for external PSU use | Build 1.3.0 and App: 0.2.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
17 | closed | USB Camera Failure ¦ Connect/Disconnect | After connecting/disconnecting USB camera more than 125 time, the device does not detect the camera. The device does not recover until we give a hard reset | High | Hardware | In Progress |  | poornesh.g;#21 |  |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | No
18 | closed | Delayed Battery Icon | Upon power up, the battery charge icon takes time to load at least 3 seconds. Before it, the icon shows red icon which show battery not connected. | Medium | Application | Resolved | 1.0.8 | davis.packiaraj;#18 | Until we read the battery status, what shall be the default state of battery? Can we hide the battery icon until we receive the status? Ayush: Hide the battery icon till you receive the status | 0.3.1 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
19 | closed | Battery Status | In on of the Hardware(Ayush has in dehli) the battery icon show charging but it is not increasing. Also, if removing the power adapter the device turn off. | Medium | Hardware | Resolved | 1.0.8-rc09 | Ayush Pareek;#16 | In current hardware it seems working. Not fully tested but we can see the charge and discharge status. Also, the battery percentage is increase and decreasing as per the defined function. Closing this ticket now, and may open when final testing will be done |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
20 | closed | FAN Control | Define the parameters to control the FAN. on what temperature the FAN should start and stop. | Low | Application | Resolved | v2.1.9 | davis.packiaraj;#18 | The FAN control should be done via application instead of the shell script. The work here to bd done is the we need to run FAN with different duty cycle as per battery and CPU temperature. 23Feb26: Davis - Implemented Fan control using Battery and Cpu temperature with Charging states. Linear interpolation used for speed control. <WW 01Jul2026>: closed due to duplicate bug#218 | App: v0.4.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
21 | closed | Audio controls | When we adjust the audio 'alert' control, it seems that the audio tones are being changed instead? | High | Application | Resolved | V0.1.9 (as shown on the GUI). latest from Poornesh on 23July2025? PROVU_V_3.2.9_APP_v.0.5.9 PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | <Ayush 02Sept2025> Fixed. Alert volume adjust the key tones and Tone volume adjust the media volume <WW 01Jul2026> : Still not fixed. When I adjust the 'alert volume', the keytone volume is changed. This is not correct. The keytone volume should be changed when the 'Tone volume' is adjusted. Davis<17Jul26>: alert volue is used previously for key press feedback - now changed it to tone volume - bug fix <WW 21Jul2026>: Retested as Admin user on v0.6.3 and not fixed. <WW 07Aug2026>: Retested in v0.6.7 and fixed. | App:0.6.7 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
22 | closed | Audio volume level quiet | The audio level for tones (keypad tones and general touch-panel feedback tones) are very quiet..? <WW 20Aug2025> : RockChip asked that we measure the voltage for the audio out. Can we measure this and make sure the voltage is set the to correct level. | High | Application | Resolved | V0.1.9 (software released on 23July2025) | poornesh.g;#21 | <05Sep WW>: Voltages were measured to be correct. Software is configured to deliver max output. Hardware needs attention to position the speaker |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
23 | closed | Broncho Image Color Changed | In recent release application, the blue color changed to red and red color to blue. where as the last application which I am using version 0.1.88 having no such issue. |  | Application | Resolved | 1.0.9 | davis.packiaraj;#18 | To reproduce this, keep Broncho Camera connected with the device. Enter into shell and execute the "reboot -f " once bootup you will see the red color changed to blue Ayush 29-07-25: it was happening because of a new flip control added and a race condition occurred. | Build 1.3.0 and Application 0.2.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
24 | closed | VL/VS cameras are not recognised | If the device is powered on (from cold) with no cameras attached, when you attach a VL or VS camera (no amp cable), to either socket, the software does not recognise the cameras. | High | Yocto | Resolved | Build 2.1.3-USB and Application: 0.3.2 | poornesh.g;#21 | Even 'v4l2-ctl --list-device' does not list the cam connected. Need to check the driver? <Ayush: 21-08-25> Tried this multiple(> 7) times, the issue does not occur. Need to confirm with different hardware before closing <WW> perform test on more device at UK <Davis: 26/08/2025> Now camera detection is based on GPIO from v0.3.4 | Build 2.1.4-USB App 0.3.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
25 | closed | Screen Capture Image Blurr | If we capture the screen while moving the camera, it captures the blurred image. | High | Application | Resolved | Build 1.3.0 and App 0.2.4 | davis.packiaraj;#18;#poornesh.g;#21 | It seems that screen data changing before capture completes. Freeze the screen until the screen capture get completed. Ayush 28Aug2025: Test by screen freeze and capture the screen. Make it for a demo to test Davis: weston-screenshooter does not work in the latest build. Poornesh to look into it. Davis 09Sep2025: Observed image blurring even with weston-screenshoot. May need to have the live stream freezing approach. Ayush170925: Stream freeze approach seems working, need review to the concept from usability point. However, if screen freeze at blur frame it will capture the blurred screen. Ayush30925: Acceptable as review by Wayne. The image blur only occur with excessive movement of camera which is not likely the use case. | App: 0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
26 | closed | USB Camera Detection Delay (200 ms) | As per discussions with the China team, it is recommended to initiate USB camera frame capture 200 ms after the camera is detected. The detection event can be identified using the corresponding I/O signal. | High | Application | Resolved | Build 2.1.3-USB and Application: 0.3.2 | davis.packiaraj;#18 | <Ayush: 21-08-25> The delay is added in the gpio detection driver file. Need to confirm the result after testing the camera plug-in/out/. This can be done only when application software integrate the IO pins for USB and Broncho camera detection. Reverting the status back to Re-open until developer confirm the application integration done for camera IO detection <Ayush: 27-Aug-25> Confirmed from davis that it is implemented | Build: 2.1.5-USB, App:0.3.5 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
27 | closed | Watchdog Reset (Software Upgrade) | In recent release watchdog reset not happening. | High | Application | Resolved | Build 1.3.0 and App 0.2.4 | davis.packiaraj;#18 | The recent implementation of watchdog should be synchronized with the software upgrade functionality |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
28 | closed | Image Control : Add RGB Colour Control | Add three more control for R, G, B color adjustment. |  | Application | Resolved |  | davis.packiaraj;#18 | This is a repeat of #530. I've removed this entry (resolved) (WW19Aug2025) |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
75 | closed | USB Camera Detection issue (From boot up) Build version 2.0.0 | In new build version 2.0.0, Application 2.7, if we plug the camera and restart the device, only one camera detected. On driver level this is fixed and we are able to see the cameras in v4l2-ctl --list-devices. | High | Application | Resolved |  |  | <Ayush: 21-08-25> This looks repeated with issue ID 220. Closing this. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
76 | closed | Software Crash | I had one instance where only one VS device was attached to the middle socket, then before the image was shown, the device was removed from the socket – the device crashed and reset. I haven’t been able to replicate this yet. | High | Application | Resolved |  | davis.packiaraj;#18 | <Ayush: 21-08-25> Tried to reproduce on the middle USB socket but can't able to see the reset. | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
77 | closed | Settings Date and Time | The date/time adjustments (from settings menu when logged in as admin account) are not implemented yet | High | Application | Resolved | Build 2.1.3-USB and Application: 0.3.2 | davis.packiaraj;#18 | <Ayush 18-08-25> The date and time can be adjusted but the date format has no effect <Ayush 21-08-25> Ask abhi or check GUI spec for more details <Ayush 25-08-25> As discussed with Abhishek 1. If date format changed by user the date display should change according to the format selected. <Ayush 26-Aug-25> Not resolved till today 26-Aug-25 <Ayush 27-Aug-25> Discuss with Abhi, it is not clearly defined from the flowchart | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
80 | closed | Wayland Screen | Some text appears on the top-right of the boot-up screen (after the splash screen) | High | Application | Resolved |  | poornesh.g;#21 | <Ayush: 18-08-25> Text is removed now | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
81 | closed | Wayland server/screen | Can we replace the grey pattern screen during boot-up (after the splash screen) with a fixed image (splash screen)? – would be good to have the progress bar continue until the home screen is shown? | High | Application | Resolved |  | poornesh.g;#21 | <Ayush 18-08-2025> The grey screen is replaced with the Blue screen | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
82 | closed | Battery Indication | With no battery attached, my unit shows 127% battery capacity and the blue led is flashing.? | High | Application | Resolved |  | davis.packiaraj;#18 | <Ayush: 18-08-25> The battery error icon comes instead of 125%. The icon is "red with ! in it" | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
83 | closed | Battery Configuration | I have a ZCF battery that does not charge. Its around 17% and discharging (if no power is attached), but there is no fault indication or other on the GUI. The GUI does not show the battery is charging when power is attached. There is defo something wrong (or different) about this battery because other batteries work correctly in this unit. I’m not sure if this is a software issue, but we need to further understand the features available to us for these batteries (both the BQ25672 and BQ40xxx devices).We can ask ZCF for help on the BQ40xxx devices. | Medium | Application | Resolved |  | poornesh.g;#21 | <Ayush09102025> : Awaiting ZCF to share updated battery settings and that to be used for battery control. Work is here to take new driver and integrate <12Jan2026>: Battery configuration registers have been defined by Flexi, Ayush will confirm the register settings. <Ayush210426>: List provided, MT need to check the configuration done successfuly. <WW 01Jul2026>: We have had no issues recently with battery charging and with charging at low or depleted batteries. Will check again in formal verification (last software tested was version 'PROVU_V_3.2.9_APP_v.0.5.9') | App:0.4.8 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
84 | closed | GUI | There is no back button in the gallery view yet | High | Application | Resolved |  | davis.packiaraj;#18 | <Ayush: 18-08-25> The "back" button implemented in the GUI | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
85 | closed | Video Recording | When I disable the microphone (from admin account setting), the audio still gets recorded during video recording. (but the symbol within the small icon at the top is correctly changed to show the mic is disabled). | High | Application | Resolved |  | davis.packiaraj;#18 | <Ayush: 18-08-25> If we turn off the microphone the audio is not recorded. | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
86 | closed | Software Crash | I’ve had several instances where the software has crashed (and reset) when I’ve started a video playback from the gallery view. | High | Application | Resolved |  | davis.packiaraj;#18 | <Ayush: 21-08-25> Not seeing reset while playing the video from gallery. closing it. | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
87 | closed | Software Crash | I had one instance where after a software crash, the thumbnails in the gallery were corrupted. Haven’t replicated this yet. | High | Application | Not able to reproduce | Build 2.1.3-USB and Application: 0.3.2 | davis.packiaraj;#18 | <Ayush: 21-08-25> Tried playing the videos from gallery and does not see any corruption. The issue is back to tester for adding details for reproduction steps |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
89 | closed | GUI Camera View | Please remove the fps number from the viewports for non-admin users. Only show the fps for admin users (for the next few weeks at least. Ultimately, we won’t show fps on the screen at all). | High | Application | Resolved |  | davis.packiaraj;#18 | <Ayush: 21-08-25> Fixed | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
90 | closed | Power Button / Sleep Mode | Sleep mode: If I push the power button while my finger is touching the screen, sometimes it does not enter sleep mode. I think it is entering sleep mode, then very quickly detecting a finger pressing the screen and waking back up. We might want to add a small time delay from pushing the power button and entering sleep mode to detecting an ‘on-press’ of the touch screen before waking up. | Medium | Application | Resolved |  | davis.packiaraj;#18;#Abhishek Shukla;#24 | Ayush09102025: In new version, if we touch screen and press power button for sleep, device goes to sleep mode. the screen activate again once we touch the screen again. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
91 | closed | Power Button / Sleep Mode | If there is a camera attached, I thought we decided to disable sleep mode (by pushing the power button)? We should check the requirements/wait for useability feedback. On the software, you can put the device into sleep mode (pressing the power button) if there is a camera attached. | Medium | Application | Resolved |  | davis.packiaraj;#18 | <05 Sep WW>: pushing this to closed because in the next sprint, the functionality of the sleep mode is changing. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
92 | closed | Notification Bar | If a 3rd camera is attached, you get the notification at the bottom status bar (which is correct behaviour). However, if the warning message auto-hides (user does not press the ‘X’), then the software does not show the yellow warning triangle in the bottom right-hand-side of the screen. If the user does cancel the pop-up warning (press the ‘X’), then you do get the yellow warning triangle. We should check this is as per requirement, but this seems strange. | High | Application | Resolved | Build 2.1.4-USB and Application: 0.3.4 | davis.packiaraj;#18 | <Ayush: 21-08-25> Not fixed <Ayush 21-08-25> Ask abhi or check GUI spec for more details <Ayush 25-08-25> The notification for three camera view defined in the GUI_v6 flowchart page 95. The current implementation should be updated as per the flowchart <Ayush 26-Aug-25> Not resolved till today <Ayush 27-Aug-25> Fixed as per GUI flowchart version 6. Extensive testing pending for all error messages which is out of scope with the dev testing |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
93 | closed | Notification | If you get a warning pop-up text (bottom status bar), then not all the text is visible (e.g. attach 3 cameras). | High | Application | Resolved | Build 2.1.3-USB and Application: 0.3.2 | davis.packiaraj;#18 | <Ayush: 21-08-25> Need to test for all warnings before closing this. For three camera attachment we can see the message <Ayush: 26-08-25> DIDO updated | Build 2.1.4-USB, App: 0.3.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
94 | closed | RTC date and time | The RTC doesn’t seem to be working or doesn’t seem to retain the setting when I initially set it on first boot. | High | Application | Resolved | Build 2.1.3-USB and Application: 0.3.2 | davis.packiaraj;#18;#poornesh.g;#21 | <Ayush: 21-08-25> Device does not retain the date/time settings <Ayush: 26-Aug-25> Tested below 1. Set RTC date and time. Power off device without the charger connected. Pass: Device retain the date-time settings 2. Remove battery, keep the charger/power adapter connected. Pass: Device retain the date-time settings | Build 2.1.4-USB and Application: 0.3.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
95 | closed | Power Button / Sleep Mode | If the touchscreen doesn’t work, you can’t switch the device off. You can try this by powering the device on (with the case open) and disconnecting the cable to the touchscreen. When you press and hold the power button for 5secs (and then release), you can’t operate the touchscreen to select ‘OFF’. We might want to consider adding a timeout, and if no press within say 30sec, if automatically turns off.? Need to discuss if there are any use-cases where this would be a problem. | Medium | Application | Resolved |  | davis.packiaraj;#18;#Abhishek Shukla;#24 | No action required. If the touchscreen is not working, then users can either just pull the battery out, or let the device drain out. The device will not be operable and will need to go back to service. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
96 | closed | displaying the software version on the USB stick BEFORE the upgrade begins | On the software upgrade screens, we should be able to see the version of the software that is on the USB stick before confirming that the version is updated. Once you agree to update the software, presently there is no means to cancel the operation. | Medium | Application | Resolved |  | davis.packiaraj;#18 | Ayush 091025: Details for command which need to read the software update version provided to poornesh. Application need integration for that. Davis11062026: This feature is implemented based on UI work flow and released in Appv0.4.9 <WW 01Jul2026> : This is resolved in version 'PROVU_V_3.2.9_APP_v.0.5.9', I can see a version of software before I decide to install. however, the version number and format differ to the verison format used seems to be for Yocto instead of the application. I have created a new issue for this. | Appv0.4.9 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
97 | closed | Auto-repeat on date/time adjustments | On the initial date/time adjustment, there is no auto-repeat implemented. | Medium | Application | Resolved |  |  |  | 0.3.1 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
98 | closed | Microphone audio is quiet | The sound volume seems ok for the informational tones, but when you record the microphone, the playback audio level is very quiet. Can this be improved? | High | Application | No Fix Required |  |  | Repeat of #570. I've closed this entry (WW19Aug2025) Ayush18092025 : Microphone no longer required |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
99 | closed | Loss of settings/gallery buttons | Loss of settings/gallery buttons: If you insert only the bronchoscope device, then click the full-view mode (so the side panel buttons disappear), if you then remove the broncho device, you get the nice picture of the system, but the settings and gallery buttons disappear. The only way out of this seems to be to re-insert a device. I think we need to have the ‘normal’ view back if the camera is removed when in full-view mode. | High | Application | Resolved |  |  |  | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
100 | closed | Deleting files in gallery view without selecting them | In gallery view, I can delete files without selecting them? I’m not really sure what is being deleted. | High | Application | Resolved |  | davis.packiaraj;#18 | Delete/Copy button will be enabled only if at least one file/folder is selected <Ayush: 21-08-25> Delete button only activates after selecting the file | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
101 | closed | Copying videos to USB stick does not work | Copying videos to USB stick does not work: The filename of the video gets copied onto the USB stick, but there is no content. | High | Application | Resolved |  | davis.packiaraj;#18 | <Ayush: 21-08-25> Video copy working | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
102 | closed | Image filenames copied to USB stick are not correct | Image filenames that are copied onto the USB stick are not in the correct format. | High | Application | Resolved | Build 2.1.3-USB and Application: 0.3.2 | davis.packiaraj;#18 | <Ayush: 21-08-25> Add more details in description like what formats acceptable. On current release, I am able to see the images in ".jpg" format. <Ayush: 25-08-25> As per DIDO the image format should be "jpg" so the this ok. However, DIDO define Image name as <CameraID>_<date>_<time>.jpg form that "CameraID" has no use as we are capturing whole screen. The DIDO need correction <Ayush: 26-08-25> DIDO updated | Build 2.1.4-USB App 0.3.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
103 | closed | Copying files to USB stick - can only copy one file at a time | I can only seem to copy one image at a time. I can’t seem to copy multiple images when I select multiple. The video file format is .h264 which is not compatible in general media players | High | Application | Resolved | Build 2.1.3-USB and Application: 0.3.2 | davis.packiaraj;#18 | <Ayush: 21-08-25> Be able to copy multiple files. However, the video file format is .h264 which is not common for windows player <Ayush: 21-08-25> Davis/poornesh need to check the conversion <Ayush: 26-Aug-25> Record video file format when copy to USB now change to .mp4. It is aligned with design inputs. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
104 | closed | Crash if connected to laptop (to USB-A port) during power-up | If I power-up with the USB cable connected to my linux laptop, it reaches the home page, then after a few seconds reboots automatically. Is this expected? | Medium | Application | Resolved |  | Wayne Williams;#13 | Issue due to conflict of host<>host and 5V conflicting. M.T have resolved on the M.T hardware by adding diode in series with output of USB 5V regulators (U40) WW to resolve and close this. (04Sep). WW added to issue trackers. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
105 | closed | Software application and Yocto build version number display | Can we find a way to version the yocto image separately to the application? As we get further into this dev, it would be good to differentiate to help debugging. | High | Application | Resolved |  | davis.packiaraj;#18 | <Ayush 18-08-25> We can see the application and build version under "Information" screen | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
106 | closed | Additional R,G,B image processing adjustment controls | Extra R,G,B controls for image processing. | High | Application | Resolved | Build 2.1.3-USB and Application: 0.3.2 | davis.packiaraj;#18 | <Ayush: 26-Aug-25> he new controls implemented for bronchoscope | Build 2.1.4-USB and Application: 0.3.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
107 | closed | GUI prompt to be shown on-press after holding power button for period of time | power button: only shows the pop-up when we release the button | High | Application | Resolved |  | davis.packiaraj;#18 | Shutdown pops up when press and hold instead of released | Build 2.1.3-USB and Application: 0.3.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
108 | closed | Top border controls are difficult to press - not centralised | Top line button press area is sometimes difficult to press - only bottom half of buttons respond - frustrating and add time delay in usability Repeat entry in 670. closed this entry (WW 19Aug) |  | Application | Resolved |  |  | <05 Sep WW>: This is due to the physical border of the display - this has been noted by D.G and will be corrected on the next batch of displays. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
109 | closed | Display/Audio settings button icon is broken | Display and audio logo broken - gear extends beyond the rectangle in the image) (04Sep) Davis and Ayush to email WW their icon. WW will ask A.G to sort out | High | Application | Resolved |  | davis.packiaraj;#18 | <Davis: 28Aug2025>: Need the correct image from Flexicare to update it. <Ayush17092025> New image provided on Monday, 15 Sept 2025 <Davis>: New icon still has the same issue. <Ayush>: Use PNG format for now <Davis>: The latest svg shared worked! It was not referenced correctly in previous release. <Ayush>: fixed | App: v0.3.8 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
110 | closed | Microphone recording level is very quiet | Mic works but cant pick up audio very well as not omni-directional and not pointed at front of display (so towards user) - very quiet audio during use as a result, not great for teaching. | High | Application | No Fix Required |  |  | Ayush18092025 Microphone no longer required |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
111 | closed | Camera flash feedback | Camera flash still needs adjusting in my opinion, its not easy to tell when a camera image has been taken. The white edge border could do with extending further into the screen. <WW 01 Oct2025>: Propose that the audio tone is changed to something more recognisable as a camera snapshot. | High | Application | Resolved |  |  | <WW 01Jul2026> : Confirmed fixed in version "Yocto_v3.2.9/App_v0.5.9" |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
112 | closed | viewports change if the image settings are adjusted | With two (or three) cameras attached, if you adjust image settings on one live view (for one camera) in dual view mode, exiting settings puts you back to just single view. This is a significant issue. | High | Application | Resolved | Yocto_v2.1.5 / App_V0.3.5 | davis.packiaraj;#18 | Ayush170925: If adjust the settings and press either "back" button or "save" button, screen goes back to dual mode. <WW 01Jul2026> : Confirmed fixed in version "Yocto_v3.2.9/App_v0.5.9" | App: 0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
113 | closed | Viewports switch sides | When adding second device or switching between two the screens swap sides. This is also a significant issue. <04Sep>: Davis reminded us that this feature was part of existing requirements - if a new camera was inserted, it will automatically be positioned in the left hand viewport This will not be fixed in this version and be changed in the next GUI version. | High | Application | Resolved | Yocto_v2.1.5 / App_V0.3.5 | davis.packiaraj;#18 | <Davis: 28Aug2025>: Can you explain in detail about the issue? The newly connected camera will always come to the left view port- as part of requirement. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
114 | closed | USB copying without a USB stick attached | You can copy files to a device without a usb being plugged in - no error warning |  | Application | Resolved | App0.3.6 |  | Ayush170925: the copy button only activates when a USB is attached. looks fixed | App: 0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
115 | closed | long press to select | Not being able to select files by long-pressing them feels counter intuitive when considering standard touchscreen UIs for the last decade. <05 Sep WW>: This was raised during useability but is not essential. Consider the effort of implementing this before confirming. <02Oct WW>: This is not essential. If it is possible, it should be implemented in addition to the existing 'select' button method because this has already been evaluated and accepted. The 'long-press' is a nice-to-have but depends on the amount of effort. |  | Application | No Fix Required |  |  |  |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
116 | closed | USB-A socket connection detection reliability | USB stick (RHS side) connection is flakey, and minor wobbles fully interrupt it (HW issue or moulding issue?) |  | Application | Resolved |  |  | Software team to check if this is caused by software or hardware. This item has also been copied to the product fault list and will be checked by hardware team. <WW 14Aug>: This seems to have been improved with recent software versions and use of 'good' USB sticks |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
117 | closed | Admin video settings text | In Admin/Video settings, the text states ‘Minimum Recorded Video Length (minutes)’. It should state ‘Maximum recorded Video Length (minutes)’ | High | Application | Resolved | App0.3.6 | davis.packiaraj;#18 | Ayush170925: text fixed | App: 0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
118 | closed | Factory reset | If we select ‘Factory Reset’, it does not wipe the admin account/password..? We were expecting it to revert back to ‘virgin’ settings such that when it resets, we should be presented with the ‘Set Language’ etc screens?. Also, we need to make sure that all media files, database entries, partion data is wiped. We had a bug in an earlier version of software that caused the software to 'think' that the remaining storage space was 0GB. We must ensure that after a factory reset, the device is cleared of data and available to use again. | High | Application | Resolved | Yocto_v2.1.5/App_v0.3.5 | davis.packiaraj;#18 | Ayush18Sept2025: This delete the error logs, video files and user accounts. However, it does not delete the swupdate folder contents After factor reset, the initial screen show "welcome, guest" where the GUI version 6 flowchart define "Welcome" only until the account configured Ayush240925: looks fixed | App: 0.3.7 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
119 | closed | USB camera detection response time | The response time of detecting a USB type camera attached seems quite long? Can this be reduced? <04Sep>: Present detection time (from insertion to camera showing) is TBD - WW will measure and report back.: measured times are as below: VS - plug in, 2.583s (socket B) VS - plug in, 2.539s (socket C) VS - plug out, 0.270s (socket B) VS - plug out, 0.320s (socket C) (Taking it out is near enough instant - comes down to how quick i can stop the watch) VL - Plug in, 2.370s (Socket B) VL - Plug in, 2.431s (Socket C) VL - plug out, 0.382s (socket B) VL - plug out, 0.220s (socket C) (same as VS for disconnecting - Comes down to my reactions) WW - I feel these insertion times are too long. Can we understand where the delays are coming from. ? We could look at providing the audio tone as soon as the software detects the GPIO signal being asserted... this will reduce the confusion to the user about if a device has been recognised. | High | Application | Resolved |  | davis.packiaraj;#18 | USB detection changed to GPIO <Ayush 02Sept2025> Need to define the detection time to test this. <Ayush 18Sept2025> Checked the DDI-000120 and does not see a clear timing for detection of USB camera <Ayush 18Sept2025> @poornesh, @davis Need discussion what minimum time can be achieved here for detection of camera. I feel as it is IO based now the detect time should reduce. <Davis 23 Sep 25>: For USB camera, After GPIO detection, it takes around 1.4 sec to get detected by the v4l2 driver. Once detected, it takes 300 ms to receive the first frame. <Ayush> Move it GUIv9, change label to backlog <Ayush 10-Oct-2025> Reduction in time delay expected <08Jan2026>: Davis is reporting approx 1.7sec insertion time but this was measured using the software and from when the GPIO was detected to when the first frame is seen. Davis will repeat this test but using a different method using a manual stop-watch and timing between when the USB is plugged in to when the image is seen on the screen. Davis-March5th - On 02/27 shared analysis report. Next steps? <01Jul2026>: Tested on software version 'PROVU_V_3.2.9_APP_v.0.5.9' -> for USB camera recognition time (from physical insertion to seeming image on the display) is now about 1.8sec. This is faster than 8" device (about 2.1sec). Although no requirement on this yet, this is acceptable for launch. | App:0.3.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
120 | closed | Top border controls layout | The top border and icons seems too close to the top of the display. Can you check the vertical position of the frame-buffer? As an example – if you capture some video, then play-back the video in the gallery, you will see the icons in the top border are more central (vertically) within the blue bar. But this is not how it looks on the home screen? |  | Application | Resolved |  |  | <05 Sep WW>: issue with the physical border on the display. Will be corrected in next samples. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
121 | closed | 3rd camera notification pop-up | If I have 3 cameras attached but only 2 selected, then tap the 3rd camera to select, there is no notification pop-up to indicate that only 2 cameras can be selected at a time.? (might need a GUI-spec update if this is not already captured) | High | Application | Resolved |  |  | <Ayush 02Sept2025> Fixed. when we tap third camera selection, pop-up appears as expected <Ayush18092025> There is one case remaining. if we tap for single view for one camera and then connect other two camera one by one, device does not show error/notification pop-up on screen. The error message does not appear even if we switch to dual camera mode <Ayush250925> Connecting third camera show error pop-up | v0.3.7 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
122 | closed | Keyboard complexity | On the keyboard – we need to try to reduce its complexity. Remove the keys for: a) ‘smileys’ b) The ‘wiggle’ pad (freeform?) | High | Application | Resolved |  |  | <08Jan2026>: Davis advised that there are more keyboard configurations available under QtV6x. The plan is to assess the impact of updating to QTV6.x and implement these keyboard adjustments at that point <Davis: 05Feb26> freeform can be eliminated, smiley / language selection can not be done in Qt5 Davis: 03/5-- Language selection+ smiley key+ scribble all were taken care through Yacto and updated in v0.4.5 & v0.4.7 <01Jul2026> : WW tested on version 'PROVU_V_3.2.9_APP_v.0.5.9' is ok. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
123 | closed | Keyboard remains in focus | If you are on the ‘MyProfile’ screen and entering some information with the keyboard, then keep the keyboard in focus and press the ‘back’ button, the GUI changes screen, but the keyboard stays in focus instead of hiding. | High | Application | Resolved |  |  | <Ayush 02Sept2025>: Keyboard hides on press of back button |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
124 | closed | Video playback controls - missing icon | When viewing a video (in gallery view), the button (Stop) at the end of the progress bar is inactive (small square and greyed out) | High | Application | Resolved |  |  | <Davis: 28Aug2025>: This icon is the Full screen Icon, which takes the video to play in full screen. | App0.3.5 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
125 | CurrentPlan | camera image processing latency | Latency measurements: Broncho + VS camera device attached: VS is 180ms Broncho + VS camera: Broncho is 160-180ms VS only plugged in : 140ms Broncho only plugged in : 150ms Test scenario is two cameras attached - Broncho + VL; recording enabled; image processing components enabled (Brightness, contrast, saturation, gamma, R, B). | High | Application | Ready for testing | Yocto_v2.1.5 / App_V0.3.5 | davis.packiaraj;#18 | Davis: We tested the cameras here with the latest application and found the latency to be ~130 ms. What is the test setup when we measured 180 ms? Ayush18092025: Need to test with the Hi-speed camera from flexicare <12Jan2026>: Sunny has completed LED/sensor based test method - results need to be analysed Davis-03/5-- Latency analysis report shared to FC team, Approx 140msec when hue and color temperature components are disabled Davis<23-Apr-26> Latency testing App released on 4/21, V5.RC01, RC02 Davis <11062026> Released in App v5.4, working on stuttering new bug Davis<25062026> Test Application w.r.t Latency single view to be released today , Wayne shall test and confirm whether we can go ahead with this approach. If we are going with this approach the timeline to implement is 10days with basic testing Davis<30062026> Cuurently as per discussion dated 6/26, This Bug is on Hold Davis<22072026> Post all usability bugs closure to be taken up, WW tested in 5 of the device in China , Latency seems abit okay, but still fine tunning is required Davis<05082026> using v4l2 to feed broncho watchdog was causing delay - changed to use ioctl | App v 0.6.8 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | Yes | ok-so far so good-re-test, v6.8 for usability test with dual view , for final product launch go for single view | v6.8 for usability test with dual view , for final product launch go for single view | 2026-08-07
126 | closed | video playback controls obscuring view | When playing back video in gallery mode, the play icon at the centre of the video is distracting. If a user wants to pause and view the detail, it is obscured by the play icon. We should consider removing the play icon – Need to confirm with Dan S and Alistair and Abhi. |  | Application | Resolved |  |  | Davis<05Feb26>: Hiden the play icon. | App:v0.4.5 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
132 | closed | software crash when trying to view the next media file | Reboots when you go past the last stored image (Bug) , a) Click on Galley b) Select and Image c) Press the &gt;&gt; until you reach the end but keep going. i) Result: Crash and reboot ii) NOTE: Was repeatable but did not happen once the device was reset | Low | Application | Resolved |  |  | Ayush 28Aug2025: Tested on new version Build v2.1.5, App:0.3.5 and does not see reset. In new software once we reach to end the forward button disappeared. | Build 2.1.5-USB App 0.3.5 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
133 | closed | RTC does not retain its value | Wrong date on the device (RTC Issue (BUG)) |  | Application | Resolved |  |  | Closed this as this is a repeat of #410 |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
134 | closed | Viewport not shown correctly | Dual camera config returns to a single camera (Bug) a) With 2 cameras connected and viewing dual screens, press one of the config buttons b) Save the config c) The device returns to the main screen but only shows the screen that was edited, not dual screens |  | Application | Resolved |  |  | Repeat of entry #590. I've closed this entry (WW19Aug2025) |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
137 | closed | Keyboard does not auto-hide | Keyboard persists (Bug) a) In the Admin profile page, press edit b) Without editing anything, press back c) Keyboard persists |  | Application | Resolved |  |  |  | 0.3.1 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
138 | closed | Back button missing in gallery view | When looking at a Video or Picture, it’s not obvious how you go back to the file library view (Observation) |  | Application | Resolved |  |  | Back button added in Image/Video View page | 0.3.1 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
139 | closed | HDMI in sleep mode | HDMI Video when the Display is in sleep mode (Inconsistent UI) a) With the display connected via HDMI, press the power button to put the device in to sleep mode b) Display goes off c) Connected HDMI display still shows a live image | High | Application | Resolved | Yocto_v2.1.5 / App_V0.3.5 | davis.packiaraj;#18 | <WW 01Jul2026> : Confirmed fixed in version "Yocto_v3.2.9/App_v0.5.9" | App V0.3.5 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
140 | closed | HDMI output different during power-up | Power up with HDMI Connected (Inconsistent UI) a) Different screen shown between the display and HDMI screen | Medium | Application | Resolved | Yocto_v2.1.5 / App_V0.3.5 | davis.packiaraj;#18 | Ayush170925: Require discussion. Ayush: Next version poornesh fixed it. Ayush250925: Consistent GUI shown in my case. Note that the HDMI monitor scanning should be faster to see the boot screen, In my monitor I noticed that it take 20 sec to scan and by that time device get booted. One thing, I noticed here that if I press power button for sleep mode when adapter plugged in, the HDMI goes off but device screen still on Reopen for further comments Ayush: If device connected with power adapter, and we try to sleep mode the HDMI goes off but the screen is running as is. Davis<05Feb2025>: When tapping powerbutton, sleep popup will be raised and device display shall also turn off along with hdmi, when put to sleep <WW 01Jul2026> : resolved in version 'PROVU_V_3.2.9_APP_v.0.5.9' | App: v0.4.5 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
141 | closed | Power button for switching OFF | Turn off the display (Observation) a) A long press off the power button does not bring up the power down option menu until the power button is released. I was holding down the power button, with a generic 3 second rule in mind but nothing happened until I released the key |  | Application | Resolved |  |  | Shutdown popup comes up when power button pressed down | 0.3.1 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
143 | closed | Continues recording in sleep mode | If you are recording video, then put the device into sleep mode (single press of the power btn), it still records video even when its in sleep mode. Requirement should be to stop recording if its in sleep mode | High | Application | Resolved | V2.0.0 | davis.packiaraj;#18 | Ayush170925: Fixed, recording stopped if we press the power button for sleep mode. In case of shutdown mode, device saves the ongoing recording but the file play back does not work for the saved recording. Ayush250925: Observations are 1. if i press the power button during recording and adapter is connected with the device, the device does not go to sleep mode but recording get stopped. This is unacceptable as power button turning of the recording 2. If recording is on and device take 4-5 second to enter into shutdown mode. This not acceaptable. <Davis>: If the recording has to be saved automatically before shutdown, the the extra time will be required. Otherwise the file may not be saved properly. <Ayush> should be added a progress bar in GUI v9 30Oct2025: Meeting Discussion - GUI Screens to be designed and to be picked up in SRS to move forward <WW 01Jul2026>: Resolved in version 'PROVU_V_3.2.9_APP_v.0.5.9' | App: 0.4.5 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
144 | closed | Search function is not intuitive | the search function in the gallery view needs some attention. Flexicare needs to define the requirements more clearly for this view. Could there be a grey text background in the search text bar that prompts the user for the date format for searching? | Medium | Application | Resolved |  |  | 30Oct2025: Meeting Discussion - Davis: Need watermark on the search bar with the required format to search | App: 0.4.1 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
145 | closed | Gallery view header text whitespace | In the gallery view, the header text should be corrected to include white space between the number of folders and the word 'Folders'. e.g. '3Folders ....' - should be a space between '3' and 'Folders'. |  | Application | Resolved |  | davis.packiaraj;#18 |  | 0.3.1 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
146 | closed | Crash behaviour when Broncho connection is intermitent | sometimes, if the bronchoscope insertion is not performed quickly and the connection is intermittent, I have seen the software crash. scenario is to very slowly insert the bronchoscope and as it is being inserted, 'wiggle' the connector to force an intermittent contact. You should see the LED flash on/off as the connection is being made. | Low | Application | Resolved |  |  | Ayush 28Aug2025: Tested on new version Build v2.1.5-USB, App:0.3.5 and does not see reset. I feel this get fixed as the detection method changed to IO pin | Build 2.1.5-USB App 0.3.5 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
147 | closed | 'LIVE' recording shown on playback | Due to the whole screen being captured during recording, the word 'LIVE' is shown during playback. This conflicts with our requirement of only showing 'LIVE' during the real-time live image streaming. Consider options such as: a. only capturing the viewport screen area during recording and adding the date/time as a layered text b. block out a portion of the top bar during recording c. TBD Note that this is a change of requirements to the present GUI (v6) |  | Application | Resolved |  |  | <05 Sep WW>: The word 'PLAY' has been overlayed ontop of 'LIVE' for the recorded video. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
148 | CurrentPlan | Watchdog reset feature improvements | The watchdog will reset if the media-capture or other non-image related processes are taking too much time (failure condition). It would be better if instead of resetting the whole system, the software attempts to keep the main camera capture and image-processing process going and only reset the other processes. this will be a change of requirements and needs more consideration before implementing. | High | Application | Ready for testing |  | ayush | 30Oct2025: Meeting Discussion - Davis: For non critical errors/ non critical process, do not reset the watchdog. Just show the warning message. | App: v0.4.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA | No change required | NA
149 | closed | Time Setting Wrap-Around | Useability feedback is that we should allow wrap-around on controls that have a wide range (date/time) | Medium | Application | Resolved | V0.3.1 PROVU_V_3.2.9_APP_v.0.5.9 |  | Davis: Date and Time will wrap around <WW 01Jul2026>: Tested on version 'PROVU_V_3.2.9_APP_v.0.5.9'. Month cannot wrap around (go from Jan (01) to Dec (12) ). Day, hours and mins can wrap. We need month to also wrap. <Davis 17/Jul/26>: Added wrapping in month - v0.6.3 <WW 21Jul2026>: Resolved in v0.6.3 | App: v0.6.3 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
150 | closed | Multiple Icon Selection Issue | Tapping the second icon does not deselect the first one. Device shows both icons selected simultaneously. Expected: Tapping a new icon should deselect the previous one. <04Sep>: Example on language screen - user can select more than one language. Also for other screens that need to have 'exclusive' selections. Abhi to add the detail to the GUI-Spec to define which screens are single selection only and which screens are multiple selection acceptable. (push for GUIv9) | Medium | Application | Resolved | 0.3.1 | Abhishek Shukla;#24 | <Davis: 28Aug2025>: Ayush, Abhiskek, can you please describe the issue more in detail. <Ayush091025> Abhi need to add details and share with MT <30Jan2026>:Defined in GUI as The GUI shall allow user to select one options at a time on the screens. (Single Touch) The GUI shall Highlight the Button colour with bule when pressed and Hold and perform action when released. Davis-03/5- This is part of touch diplay handling and no fix is required . <WW 01Jul2026>: accepted requirement and closed for version 'PROVU_V_3.2.9_APP_v.0.5.9' |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
151 | closed | No Red Border on Empty Fields | Account creation page does not highlight empty text boxes with a red border as expected when left empty.(V6 dev) | Medium | Application | Resolved | 0.3.1 | davis.packiaraj;#18 | Davis: Not clear on the expected behaviour. 12Sep2025: This is not part of GUI6. Need to move it to next version. Ayush: 091025: Details for the red borders should be added in GUI v9 <WW 01Jul2026>: Resolved in version 'PROVU_V_3.2.9_APP_v.0.5.9' | App: 0.4.4 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
152 | CurrentPlan | Text Styling Mismatch | Text size and style do not match the GUI design. Text appears bold in the GUI but not on the device.(V6 Dev) <04Sep>: WW to discuss with A.G about any missing font packs. Davis said he need other font packs for bold extensions if we need to implement the bold feature of the font | Medium | Application | In progress | 0.3.1 PROVU_V_3.3.3_APP_v.0.6.3 | Wayne Williams;#13;#davis.packiaraj;#18 | <02Sep2025>Davis: Can we get the font pack to be used as per Wireframe? Ayush170925: Font file provided on Monday 15 sept 2025 Davis 230925: Direct loading of font file didnt work as expected. May need to move this to next GUI version. Ayush091025: The font files already provided. The target for implementation is v9 Davis <05Feb2026>: Font file provided does not match the styling as UI/UX. Abhi to confirm if we are using the correct font file, or continue using the old default font if it is acceptable. Davis: New font received and applied to application, which matches the UI/UX <WW 22Jul2026>: New font packs have been given to M.T (including bold) - (sent approx 9th Feb2026). Tested on v0.6.3 and only some text have been implemented as bold. All bold fonts to be implemented as per GUI design. Davis<09Sep2026> To be implemented in v7.6 Davis<24092026>: 50% done , 100% to be released in October version release Phase-2 | v0.7.7 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA | Check and resolve as per GTD-989 | 2026-08-25 | 2026-08-26
153 | closed | Special Character Restriction Missing | Special characters can be entered in fields like Forename, Surname, and Username, contrary to expected behaviour. Device should exclude special characters in user data input. | Medium | Application | Resolved | V0.3.1 | davis.packiaraj;#18 | Davis: This may not be feasible with this version of QT Virtual keyboard 11Sep2025: Discussed in meeting to not restrict in input, but to add validation on username, forename and surname. Ayush170925: The text strings for error not correct. For example if enter special character in forename, the error message says "firstname, forename, should not have special character" Ayush250925: Error message corrected | App: 0.3.7 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
154 | closed | Password Policy Not Enforced | Device does not enforce the password policy: at least one capital letter is not mandatory. | Medium | Application | Resolved | 0.3.1 | davis.packiaraj;#18 | Ayush170925: Show error for upper case letter, small case letter and a number. however, if miss to enter a special character, account creation failed without any error Davis 250925: Hashed password saved as blob in database. Ayush250925: The main query here is that, if I miss to enter the special character then it should give error. Also, it should raise error for common passwords like 1111, 12345 <Davis> Ayush, I dont think Special character is part of password policy. There is a list of common password provided by Flexicare, which is already enforced. The UI workflow only shows restriction of having atleast 1 small char, 1 capital char and 1 number. Let us know how to proceed on this. <Ayush> Acceptable. As the device allow user to enter special character | App: 0.3.9 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
155 | closed | Incorrect Navigation After Reboot | After the second reboot, the device navigates to the Home page instead of the Sign-in page. Likely due to sign-in screen not being added at boot. <04Sep>: Abhi to update the GUI-spec and workflow. The present software implementation is correct -after reboot (and when the device has been configured after first-time use), it should boot straight into the home screen (camera views) | Medium | Application | Resolved | V0.3.1 | Abhishek Shukla;#24 | No further action required |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
156 | closed | Welcome Screen/ Account success page Graphics Mismatch | Welcome screen / Account success page Graphics do not match the GUI design. | Medium | Application | Resolved | V0.3.1 | davis.packiaraj;#18 | Fixed, using latest icon. Ayush170925: On top message of welcome <username> "!" mark missing Ayush250925: Resolved | App:0.3.7 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
157 | closed | Home Page Icon Mismatch | USB and HDMI connection icons on the home page do not align with the expected GUI workflow/design. Also the image of display not matching with GUI. <04Spe>: The image on the present software is correct, but the GUI-spec image is not correct - the USB-A image and HDMI image need to be swapped. Abhi will update GUI-spec in v.9 | Medium | Application | Resolved | v0.3.1 | Abhishek Shukla;#24 | Davis: Icons are not in correct sequence in the UI/UX. The icons are displayed as how the physical port are located in display. Display Icon is received from Flexicare. There seems to be some icon rendering issue, Get an updated Icon. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
158 | closed | Inconsistent Time Format | On first boot, time is set in 24-hour format, but Home page displays it in 12-hour format. <05 Sep WW>: Proposal is to fix the format to MM-DD-YYYY and use 24hr format everywhere | High | Application | Resolved | V0.3.1 | Abhishek Shukla;#24 | Need to fix RTC display time format on top of the page 30Oct2025 - Meeting Discussion - Davis: - Do not use MM/DD/YY or DD/MM/YY selection in settings screen. Fix format to MM-DD-YYYY and use 24 hours clock throughout. Davis: v9 UI uses dd MMM yyyy format for display and mm dd yyyy for settings, and 24 hr clock, which is followed throughout <WW 01Jul2026> : Time format Resolved in version 'PROVU_V_3.2.9_APP_v.0.5.9' - uses 24hr in adjust screen and home screen. | App: 0.4.1 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
159 | closed | Bronchoscope Icon Size Missing | Bronchoscope connection icon does not show expected size and label information. | Low | Application | Resolved | V0.3.1 | davis.packiaraj;#18 | The icon is changes in new gui v9 flowchart. so this not applicable |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
160 | closed | Unsaved Settings Not Reverted | On the image properties settings page, pressing the Back button without saving still navigates to the main page but does not revert changes <04Sep>: WW to discuss this with the team. The use of 'Revert', 'Back' is unclear and there is no button to change the settings back to the 'default' (out-of-box) settings. | High | Application | Not able to reproduce | V0.3.1 | Wayne Williams;#13 | <Davis: 28Aug2025>: Is this not the expected behaviour? <Ayush091025> This becomes invalid as v9 flowchart has updated the controls behaviour |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
161 | closed | Incorrect Button Colors After Pressing | After pressing, the Back, Revert, and Image Properties buttons do not match the expected colour behaviour as per GUI. | High | Application | Not able to reproduce | v0.3.1 | davis.packiaraj;#18 | Ayush170925: Look for "back button graphic still not as per v6flow chart. "Back" button should be light blue. Ayush to confirm from Andrew G Ayush230925: Mail sent to Andrew G Ayush: move it to next version 9 Ayush: Version 9 has the changes for the control buttons. Resolving this as it should be covered by flowchart with new changes |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
162 | closed | Missing Title on image properties Settings Page | The label “Image Output Settings” is missing from the Image Properties Settings screen. | High | Application | Resolved | V0.3.1 | davis.packiaraj;#18 | Fixed. | App:0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
163 | closed | Limit Button Colors Missing | The '+' and '–' buttons for maximum/minimum limits do not change color as per GUI design. | Medium | Application | Resolved | v0.3.1 | davis.packiaraj;#18 | Buttons colour shown as blue when we press. As discussed with Abhi, if we reach to min or max the color should be white. However it is not clear with the Version 6 flow chart <WW 22Jul2026>: zoom +/- buttons are not having the same colour as per GUI spec (GTD*989) <WW 24Jul2026>: ReTested in v0.6.4 and resolved. | App:0.4.1 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
164 | closed | Swap Button Not Updated | Swap button still appears as per the old GUI. Additionally, it does not change to white background after tapping as expected in GUI V6. |  | Application | Not able to reproduce | v0.3.1 | davis.packiaraj;#18 | This should be referred from the new flowchart version 9 as it has modification in this control button |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
165 | closed | Camera Icon Color Mismatch | After connecting the third camera, the icon color does not match the expected GUI behavior. |  | Application | Not able to reproduce | V0.3.1 | davis.packiaraj;#18 | The camera connection icons are changed in the new flow chart version 9. Refer the new flowchart. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
166 | closed | Gallery/Settings Buttons Non-Functional in Subfolders | While inside a subfolder, tapping the Gallery or Settings button does not perform any action. Expected: Should return to main Gallery or Settings screen. | Medium | Application | Resolved | v0.3.1 | davis.packiaraj;#18 | <WW 01Jul2026> : Confirmed fixed in version "Yocto_v3.2.9/App_v0.5.9" | App: 0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
167 | closed | Incorrect Camera Feed Reinitialization | Two cameras connected ON port A and C. connect 3rd camera to Port B. keep on only A and C camera feeds, tap on the gallery button. Go back to the home screen, device shows all 3-camera connected with all three-camera icon blue, feeds only as previously A and C | High | Application | Resolved | v0.3.1 | davis.packiaraj;#18 | Ayush170925: Looks fixed | App: 0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
168 | closed | Gallery Scroll Bar Issues | Scroll bar in Gallery is not functioning correctly and does not match GUI design. <04Sep>: Davis showed the challenges with the standard Qt control of the scroll bar and this is not exactly matching the desired scroll style of A.G design. Davis will make the Qt control look similar to our design and email WW. Discuss with A.G if this is acceptable | Medium | Application | Resolved | v0.3.1 | Wayne Williams;#13;#davis.packiaraj;#18 | <02Sep2025>:Davis: We may not be able to exactly match the Scrollbar as in Wireframe, as scrollbar has limitations on customization. Can this be made low priority, as this may take lot of time to explore possibilities to customize the scrollbar. <11Sep2025> Davis: Discussed with Ayush on meeting. The new custom scroll bar design will be used for the release, and if color or size changes required, it will be taken up later. <Ayush: 18sep25> Scroll bar is accepted from AndrewG. | App: 0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
169 | closed | Missing File Info During Preview | File size and other details are not shown at the top during preview of videos or images. | Medium | Application | Resolved | v0.3.1 | davis.packiaraj;#18 | Ayush170925: Date and time stamp missing from the information details. screen ref is 42 on XD v6 files Ayush250925: For video file size, time, date and duration shown. For Image file size, time and date shown. This is correct as no duration applicable for image capture. | App: 0.3.7 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
170 | closed | Missing Video Duration Display | Video playback screen does not display total duration at the end of the progress bar. | Medium | Application | Resolved | v0.3.1 | davis.packiaraj;#18 | Ayush170925: fixed | App: 0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
171 | closed | No Save Confirmation | Pressing the “Save” button does save settings, but no confirmation message (e.g., 'Setting saved') is displayed. Abhi to provide more details about which screen this happens (WW 02Sep2025) | Medium | Application | Resolved | v0.3.1 | davis.packiaraj;#18 | Davis: Fixed on Display and Audio Settings Page. Ayush170925: Fixed on the Display, Audio setting page. Account create page also have save button but that is missing on the GUI version 6 flowchart Davis-03/5- Shall be taken up in next release | App: 0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
172 | closed | Event Log Page Missing Scroll Button | No scroll option available on the Event Log page. | Medium | Application | Resolved | v0.3.1 | davis.packiaraj;#18 | <11Sep2025>Davis: Will use the custom scrollbar as agreed in meeting with Ayush. Any changes if required can be taken up later. Ayush170925: Review it before closing this Ayush 180925: AndrewG accepted the scroll bar graphics shared by MT | App: 0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
173 | closed | Event Log Export Not Working | Export function does not work. After confirming export, the Event Log is not exported. | Medium | Application | Resolved | v0.3.7 PROVU_V_3.2.9_APP_v.0.5.9 PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | No Fix Required! Working as expected. Need to Re-Test. Ayush170925: File exported as csv but transfer is so quick to see the transfer screen. There is a screen in version 6 flow chart for file transfer complete. refer pages 77 and 82 as per version 6 XD file Ayush250925: The transition is too fast so the screen 77 can not be seen. Page 82 is implemented but the page view is not same as the flowchart. closing this as functionality seems ok Davis-03/5-- Released in v0.3.7. No fix required as of now. <WW 01Jul2026>: Tested on version 'PROVU_V_3.2.9_APP_v.0.5.9', a file does get created on the USB stick, but there is no data in the file (0Bytes). tried several times with a lot of events in the log on the device. <Davis>17July2026: Added stream flushing before closing file - ensures large content is written before files being closed <WW 21Jul2026>: Retested on v0.6.3 and I can now see an event log file, but the number of files in the log is much less than all the files I see on the device. I can see over 100 events on the device, but only 7 events contained in the logfile (view with excel) <WW 17Aug2026>: Retested in v0.7.0 and fixed. | App: 0.6.3 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA | Davis to check on this | 2026-08-11 | 2026-08-14
174 | closed | Cannot record video or capture images if all files are deleted | If I take some videos/images, then go into the gallery view and delete All files, this deletes them correctly. Then if I go back to home screen and try to take more images/video recording, it doesn't record the images/videos. When I go back into gallery view, there are no images/videos. I have to reset the device to fix this. | High | Application | Resolved |  | davis.packiaraj;#18 | <Ayush 02Sept2025> This is fixed. I can see photo and video | App:0.3.5 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
175 | closed | Software crash after multiple attempt on Deleting all files and does not boot after 94% | After deleting all folders with "Delete All" button and after performing multiple operations more than 2 or 3 times , device get crashed and does not reboot after 94%. <WW 14Aug>: To add to this, I also had this same problem and when I repeated this, the device is now 'bricked' It is stuck on bootloader at 94% and even if I reset multiple times with/without battery (on main power), it still stuck on this screen. I can't seem to establish ADB via USB comms. It might be that the 'active' application is getting corrupted when the file(s) are deleted from the gallery because when it crashed the first time, after the re-boot, I was presented with the initial language setup, date/time etc screens..? To fix this, I had to open the case, enter 'recovery mode' and reflash the .wic image using the linux upgrade_tool app. | High | Application | Not able to reproduce | Build 2.1.3-USB and Application: 0.3.2 | davis.packiaraj;#18;#poornesh.g;#21 | <Ayush: 21-08-25> Davis disabled the delete button until user select a file. Also, delete button or delete all will be deactivated when there is no file Need to find out the root cause for settings reset |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
176 | closed | Software update can be performed using micro-USB port (camera ports) | The micro-USB ports (camera ports) can be used for software upgrade. This should not be allowed. We would like M.T to understand how to fix this (find solution), but not implement the solution just yet because this is useful during development. | Medium | Application | Resolved | 'core-image-minimal-rockchip-provu2-20250804111126-v1.3.0-ADB.wic' | poornesh.g;#21 | Ayush170925: Tested over both micro usb port and no usb get mounted. Software update fail in case USB not detected. | Yocto version - 2.1.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
177 | closed | Not charging when PSU is attached (sometimes) | Sometimes, after using the device normally when its running only on battery (e.g. doing some recording, attaching/removing cameras etc), when I insert the external PSU, it does not detect it being attached and does not charge the battery. I need to reset the device for it to correct itself. | High | Application | No Fix Required | App: V0.3.1; Yocto build: v2.1.1-USB | poornesh.g;#21 |  |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
178 | closed | Battery Charging Detection | Ayush: 26-Aug-25: Remove the battery charger for more than a minute and connect back. Device takes more than 55 seconds to detect the charger. | High | Application | Resolved | Build: 21.4-USB, App 3.4 PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18;#poornesh.g;#21 | <Davis: 26Aug2025>: When we connect the charger, the battery state changes from 'Discharging' to 'Not Charging', and then changes its state to 'Charging' after a min or so. Is it ok to Display charging icon when the battery state changes to 'Not Charging'? or do we need to change anything on the driver?@Ayush, @Abhi please confirm. <Ayush 18Sept2025> if battery goes below 25% device shows yellow icon but it does not show charging symbol. <Ayush: 26-08-25> @poornesh need to look on control layer for the delayed response on the charging status <Ayush: 27-08-25> The problem occur because the driver report the charge status as "not charging" after plugging the charger. The staus get change from "not charging" to charging after a minute. This happens randomly. In current release, the device will show the status charging even in case of "not charging". @poornesh need to look why driver report "not charging" Davis: Using path of bq25790 to get the charging state. which is giving the state as expected. <WW 01Jul2026>: Tested on version 'PROVU_V_3.2.9_APP_v.0.5.9', acceptable times for insertion and removal. Approx 1.7sec for removal detection and <1sec for insertion detection. Both audio and GUI icons updated in this time. | App: 0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
179 | closed | Software Upgrade ¦ Watchdog Reset | Ayush 26-Aug-25: Device does not take auto reset after the software upgrade.. This was fixed previously. | High | Application | Resolved | Build: 21.4-USB, App 3.4 | davis.packiaraj;#18 | <Ayush: 27-Aug-25> Fixed | Build 2.1.5-USB App: 0.3.5 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
180 | closed | Protecting application data | We need to improve the robustness of the software to prevent the application deleting files that are essential for the application itself and prevent hanging/crashing/corruption. Suggested controls to be considered: 1. if the boot-up is not able to boot into the application, we should have some sort of feedback to the user about this (instead of being stuck on progress bar) 2. It should not be possible (or very very difficult) for the application to delete any important files that could prevent the application from booting or running normally. 3. If one of the partitions (for A/B software update) is corrupted, does it get 'corrected' on a subsequent boot-up? (e.g the non-corrupted partition is copied into the corrupted partition)? | Medium | Application | Resolved | Yocto_v2.1.5/App_v0.3.5 |  |  |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
181 | closed | Wifi and LAN not working | <Ayush 28Aug2025> I tested the wifi and Lan via terminal and found that these are not working. I used command "ip a" and doesn't see the IP address listed for eth0 Ayush will use China BuildRoot version to test the hardware. | High | Application | Resolved | Build 2.1.5-ADB App 0.3.5 | poornesh.g;#21 |  |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
184 | closed | USB cameras (VL/VS) 'Camera Error' shown sometimes when attached | Sometimes when I connect a VL or VS type camera to the display, I see a brief text of 'Camera Error / Replace ProVu input device' within the viewport before it shows the camera image | High | Application | Resolved | Yocto_v2.1.5/App_v0.3.5 | davis.packiaraj;#18 | <04Sep2025>Davis: Camera error detection timeout increased to 5 sec from 1 sec. Ayush170925: 1. if I connect a faulty camera then device shows the white screen for 3-4 sec and then shows error 2. After removing the faulty camera, if we connect the working camera, device shows the error for 3 seconds then shows the camera display which is different if I connect the "working camera" first after power up Ayush23092025: Device does not detect the faulty camera over USB ports. I can see the IO status getting change correctly when I plug a faulty camera. Davis 04Feb2026: Couldnt reproduce the issue <WW 01Jul2026>: Resolved in version 'PROVU_V_3.2.9_APP_v.0.5.9' | App:0.3.7 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
185 | closed | Incorrect error message displayed when trying to copyTo USB with no usb stick attached | When in gallery mode and viewing a video, if I select 'CopyTo' with no usb stick attached to the device, it says 'File operations failed / Not enough storage space for copying' If there is no usb stick attached, suggest either greying out the 'CopyTo' button, or provide a more intuitive warning prompt message | High | Application | Resolved | Yocto_v2.1.5/App_v0.3.5 | davis.packiaraj;#18 | <4Sep2025>Davis: 'Copy To' button is disabled if no usb detected. <Ayush 170925> fixed | App:0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
186 | closed | touch tones in the Date/Time adjustment screen are missing? | There don't seem to be touch-tones played when I adjust the date/time settings from the admin menu? | High | Application | Resolved | Yocto_v2.1.5/App_v0.3.5 | davis.packiaraj;#18 | Bug Fixed. <WW 01Jul2026>: Resolved in version 'PROVU_V_3.2.9_APP_v.0.5.9' | App:0.3.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
187 | closed | Copying files to USB stick that already has same files does not show warnings | Ayush 170925: If I try to repeat the copyTo for selected folders/files onto a USB stick (that already has the same folder/files on it), there is a very brief pop-up flash (too fast to read). can we have a confirmation screen that warns that the files on the USB stick need to be cleared before the operation? <WW 01Jul2026>: See demo video '../v0.5.9 issues/Copy To USB overwrite existing filenames'. If I try to copy individual files to the USB stick, the software checks for existing (on the stick) and gives me a warning to proceed and this is ok. But if I select a complete folder (shown in video), then it gives me the pop-up warning, but then gives me an error after that. | Low | Application | Resolved | Yocto_v2.1.5/App_v0.3.5 PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 | Ayush170925: A confirmation screen added to tell user that the action will overwrite the file. There are two option given to user "yes" or "no". If I click yes then nothing happen. Ayush18092025: The screen mentioned in version 6 XD file, 85 and 89 not completely implemented. Button colours and valuable info missing Ayush25092025: Please refer the screens carefully and implement the right colour and graphics. Ayush290925: Need to verify the file overwrite method as well. so if there are files already present with same name they should be overwritten. <WW 22Sep2026>: Tested in v0.7.7 and resolved | App v0.7.2 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA | 1. Update the SRS , if folder is copied check the files on each folders in destination for over writing | 2026-08-18 | 2026-08-26
188 | closed | Cannot update software if running only on external power | If the battery is removed, but the device is powered on external PSU, the software update should still work - but presently it blocks this: "Cannot update software" / "Battery is too low. Please connect to power or charge to atleast 25%". also, check spelling on 'atleast' | High | Application | Resolved | App_v0.3.7 | davis.packiaraj;#18 | Fixed Ayush 17092025: 1. If battery removed when external PSU connected battery error icon appears. The software update work 2. If battery is less than 25% then software update does not work even the PSU connected. Refer Design Input PR56.05 Ayush250925: 1. If connect a battery less than 25% and then try to upgrade software without the power adapter, it gives me error which is ok. 2. Once error occur, I press the "back" button and connect the power adapter and try to upgrade same error repeats. The error only clears If I go to home screen Ayush260925: May change the requirement for the error page in new gui version 9. so pushing this to next version Davis-03/5- Updated and released in v0.3.7 <WW 01Jul2026> : Tested on version 'PROVU_V_3.2.9_APP_v.0.5.9' and I can now still upgrade the software even if battery is removed. More testing to be done in formal verification, but closed for this particular issue. | App:0.3.7 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
190 | closed | Updated Bronchoscope camera image processing settings are required | We need to update the default image settings for the bronchoscope camera as follows: Bronchoscope: Gamma : 53 Brightness: 0 Contrast : 9 Saturation: 15 Sharpness: 0 Hue: -6 Colour Temperature: 2 Red: 15 Green: Disable Blue: Disable Application image processing Pipeline priorities: colour convert: 1 horizontal flip: 2 black level: disabled brightness/contrast: 3 Gamma: 5 DeNoise: Disabled CLAHE: Disabled Saturation: 4 Sharpness: disabled Hue: 7 Colour Temperature: 8 Red: 9 Green: Disabled Blue: Disabled | High | Application | Resolved | Yocto-V2.1.5 / App_V0.3.5 | Ayush Pareek;#16 | <12Jan2026>: Latest software does contain the Bronchoscope settings but need to be tested. Davis<11062026> Latest settings shared by AYush in email updated in App v0.5.5 <WW 16June2026>: These settings have now been updated again and are being managed by the SRS content. | App v0.5.5 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
191 | closed | image settings adjust controls | For the image settings adjustment, in the production release of software we should do the following: a. only allow the following controls to be accessible to the standard user: - brightness - contrast - saturation - Colour (using Red/Blue Look up table) b. we need to remove the numbers associated with each slider c. we need to review the limits for each of the controls. I don't think the values need to be as 'extreme' and we can reduce each min/max level. This will also make it easier to do fine control. WW still needs to define the LUT for R/B adjustment (colour) | Medium | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 | davis.packiaraj;#18 | Davis<25062026> Flexicare to share the details of lookup table( limits), expected by 6/29 Davis<30062026> Details received by FC on 6/26, To be released in 6/30 v6.0 <WW 16Jul2026>: Tested ok with v0.6.2 | App v0.6.0 |  | Yes
192 | closed | Zoom button locations | Feedback from formative useability study - the zoom + and - buttons are the wrong way around. | High | Application | Resolved | Yocto_2.1.5; App_0.3.5 |  |  |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
193 | closed | Field view when keyboard is in focus | For some screens (such as initial username/password entry), the keyboard obscures the edit boxes. Formative useability feedback showed that users were not able to easily understand that more data is required to be entered in the edit boxes hidden by the keyboard (e.g. password confirm). suggest either: a. placing all the edit boxes within a scroll group so that the user can access all boxes while the keyboard is still in focus b. limit the placement of all edit boxes to the window area above the keyboard (when in focus) and add more screens on that the user can navigate forward/backward to. c. other? | Medium | Application | Resolved | Yocto_2.1.5; App_0.3.5 |  | 30Oct2025: Meeting Discussion - Davis: Use option b. for now and if doesn't work look for other option. 05Feb2026: A dedicated text box has been added at the top of the virtual keyboard, which becomes visible whenever the keyboard is active. Any text entered through the input field is mirrored in this area, ensuring that the typed content remains visible to the user at all times. | App: v0.4.5 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
194 | closed | Three Camera View and Error Popup | Ayush18092025: If we press one camera for single view then connect other two camera one by one, device does not show the error message that only two camera can be seen at a time. The error does not appear even if we switch to two camera view | Low | Application | Resolved | build 2.1.6 and App 0.3.5 | davis.packiaraj;#18 | Ayush300925: The layout for camera view will change in GUI v9. Moving this to next version |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
195 | closed | Image control button freeze sometime | Ayush180925: if we tap image adjustment setting button, sometime it remains blue and does not recover back | Low | Application | Resolved | Build 2.16 and App 0.2.6 | davis.packiaraj;#18 | Ayush240925: Move it for next GUI version as there are change regarding this area. Ayush101025: Closing this as the new GUI version 9 will modify the current flow. So this should be tested with new implementation. |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
196 | closed | Hour Minute Button Roll over | Ayush 180925: Hour, Min button get roll over. for min if we keep pressing "+" button it will roll back to 0 Where if we do press "+" button for day field it does not roll back to 1 | Low | Application | Not able to reproduce | Build 0.2.6 and App 0.3.6 | davis.packiaraj;#18 | Usability feedback allows roll-over. closing this for now. refer version 9 for latest updates and clarity |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
197 | closed | Date Time Change | Ayush18092025: If we configure the date/time and the record a file or capture screen log. device take older date. Device also take 04-Aug-2017 date until user configure the date/time. it can be 01-Jan-2025 if no date/time configured | Low | Application | Resolved | Build 0.2.6, App:0.3.6 | davis.packiaraj;#18 | Davis: Folder will be created for the current date after date change. Date 04-Aug-2017 is default from RTC, this may happen if device does not have battery after setup is done. Ayush250925: fixed | App: 0.3.7 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
198 | closed | Software Reset | Ayush25092025 1. Export the error log with connected the USB driver. This will raise an error for no usb mounted 2. On home screen, start recording and press the setting button. This will raise error for "stop recording to enter into settings .." 3. Tap on galley button and you will the ProVu application crashed | Low | Application | Resolved | Build 2.1.7 and App 3.7 | davis.packiaraj;#18 | Ayush3092025: Software crash is resolved. However when there is no usb device connected the screen ID 72 (Adobe XD file) should appear which is not happening. Carefully observe the flowchart and implement the required screen. Davis 20092025: Ayush, The reference document for GUI v6 changes from v2 is from the link given below, and the expectations is to make the screen changes. As per our understanding and initial discussions with Wyane, we have concluded to work on these 30 items outlined on the document. We have met that expectations. If there is any new addition screens/requirements related workflow, we shall take it up in the upcoming version. "https://flexicaremedicallimited.sharepoint.com/sites/ProVu2MindteckShare/Shared%20Documents/Forms/AllItems.aspx?csf=1&web=1&e=EHcd89&CID=1ff6d0f8%2De613%2D4a8f%2D8a21%2D446d2e0b7d73&FolderCTID=0x0120008656A1760815C842B06DE6DAE3DD92E5&id=%2Fsites%2FProVu2MindteckShare%2FShared%20Documents%2FGeneral%2FAdditional%20GUI%20Screens%2017JUN2025" Ayush 101025: As discussed the flow chart provide by Andrew Gordon should be covered. This should be considered while reviewing the flowchart for new work package <WW 01Jul2026>: Tested on version 'PROVU_V_3.2.9_APP_v.0.5.9' and seeme resolved. Does not crash when I tested that scenario and provides a warning pop-up if user tries to export log when there is no USB stick attached | App: 0.4.6 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
200 | closed | Software Update Screens | The update screens does not align with the defined screens in GUI version 6 flowchart. This is discussed that it will be implemented in version 9 work package changes | Low | Application | Resolved | Build 2.1.9 and App 0.4.0 |  |  |  | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA
201 | CurrentPlan | Requirement that needs to be added in SRS | The GUI shall follow the password guideline defined in Flexicare Cyber (GTD-000759) - use password that contains at least 8 characters- not use password define in common blacklist Ex: 0000, 1111 GUI shall allow username and password case sensitive GUI shall allow unique usernames can only be used. GUI shall exclude special characters and keystrokes at the user data entries. GUI must allow passwords to include at least one capital letter | Low | Multiple | Ready for testing | PROVU_V_3.2.9_APP_v.0.5.9 | Poornesh/Devi | Davis-03/5- Added in SRS and shared to FC team. However Mindteck to capture in correct SRS template <WW 01Jul2026>: mostly ok but I can enter a 'smiley' character from the keyboard. I know MT have removed most of the 'special' characters from the keyboard, but some still remain. | App:v7.0 | sites/ProVu2MindteckShare/Lists/ProVu2 Large Display Software Bugs Fixes Record | NA | Check and resolve -remove the smiley key | 2026-08-14 | 2026-08-17
202 | closed | OV428 settings sent from RockChip clashing with FPGA transmission | On the latest software (PROVU2_V_2_1_9_APP_0_4_0) if I connect a bronchscope, the image is very dark. | High | Application | Resolved | Build 2.1.9 and App 0.4.0 |  | It appears that the RK software is sending the OV428 configuraiton settings too quickly (on insertion of the bronchoscope) and this is clashing with the FPGA transmission. The result is that the OV428 is not being configured correctly. <12Jan2026>: The feature needs to be robust. presently using a delay before sending the OV428 settings (after the FPGA) is not reliable. Software needs to check for first frame from OV428/FPGA before sending the correct settings. Software also needs to confirm (read-back) the settings directly from the OV428 (I2C is now connectred to RK) - but this can be done in parallel. | App: v0.4.4 |  | NA
203 | closed | Software update issues with incorrect filenames | If I have an incorrect filename on the USB stick - e.g. 'update-bundle-rockchip-provu2 (5).raucb', the software allows me to start a software update; it will search the USB stick for updates, then it will show the next screen for downloading, but the progress bar gets stuck on 0%. There is no prompt, or notification that there ar eno valid files on the USB stick. It is possible to touch the screen during this state and install software - but software does not get installed. | Medium | Application | Resolved | Build 2.1.7 and App 3.7 PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 | Davis<23-Apr-26> Taken up in SRS changes Milestone 4 -Req ID_29 <WW 01Jul2026>: Tested on version 'PROVU_V_3.2.9_APP_v.0.5.9'. : I tested again with no suitable upgrade package on the USB stick. The software allows me to proceed to downloading the software... but there is nothing to download. So I'm not sure what it is downloading. It does time-out with a 'System Error', but this is not intuitive. It should say something like 'no software found on USB drive'. After this 'system error' pop-up, I can then go back into software upgrade and proceed, and then I get the screen 'Would you like to update the current software' shown. then the cycle (download -> system error) is repeated. See video '../v0.5.9 issues/sw upgrade no USB.mp4' for example Davis<26Aug2026> 'No Software Found' GUI screen not available. Error screen will popup if no software found after the scanning progress. <WW 22Sep2026>: Tested in v0.7.7 and resolved | App v0.7.2 |  | NA | Update GTD with screen spec, Davis to implement the same.
204 | closed | EFT / EMC immunity recover issues for USB device | During electronic fast transients / EDS / EMC immunity testing, it is possible for noise to corrupt the isolated camera ports which results in a variety of symptoms such as: a. USB cameras become 'detached' b. Video feed from the USB cameras drops out c. Bronchoscope camera LED flashed on/off (during the injected noise test) The latest software (Build PROVU_V3_1_1_APP_0_4_2) does detect some of the USB corruption conditions but there are two issues: 1. When the software re-connects the USB camera, it increments the driver ID. This will cause an error eventually. 2. There needs to be a re-try mechanism implemented such as 5-times before it displays a camera error | High | Application | Resolved | PROVU_V3_1_1_APP_0_4_2 PROVU_V_3.3.3_APP_v.0.6.3 | Poornesh/Devi | <WW 21 Jul2026>: Ayush has the logs and diagnostic info <Poornesh/Devi>: Logs received from Ayush, Devi checked on the same, we do need some more details , Devi shall discuss with Ayush on 7/23 morning Poornesh/Devi<28Jul2026>:Unable to reproduce the bug due to the unavailability of the EFT test kit. Requested Wayne to share the logs via email. Davis<08Aug2026>: Reset the port if camera error is displayed for 5 sec or more. <WW 07Aug2026>: China team tested v0.6.7 and still failed ESD (on VL camera). <Poornesh/Devi 12 Aug 2026> Recovery mechanism added in GPIO detection driver -- As of now USB 1 ( PORT B) camera seems recovering post ESD strike , tried for 25 strikes. WIP for USB 2(PORT C) <WW 18Aug2026>: China team tested 'v3.3.9-USB' version - Poornesh release this on Friday14Aug2026. Link here : https://flexicaremedicallimited.sharepoint.com/:f:/r/sites/ProVu2MindteckShare/Shared%20Documents/General/Yocto_Migration/Release_for_EFT_EMC/Release_14_Aug_2026?d=w0524fb66ad7f46a3ad480fa5b398b46f&csf=1&web=1&e=oFTdHY Testing passed with some hardware mods (needed to add foam conductive pads to connect the rear of the display to the EMC can shields. | v3.3.9-USB |  | Yes | Not for usability testing | 2026-08-14 | 2026-08-17
205 | CurrentPlan | EFT / EMC immunity recover issues for Bronchoscope device | During electronic fast transients / EDS / EMC immunity testing, it is possible for noise to currupt the isolated camera ports which results in a variety of symptoms such as: a. USB cameras become 'detached' b. Video feed from the USB cameras drops out c. Bronchoscope camera LED flashed on/off (during the injected noise test) I am not sure yet how the latest software handles the bronchoscope input but presently there is no such failure because the video feed from the bronchoscope is still being streamed during the EMC test. However, we need to make sure there is sufficient control in the software to detect if the streaming has stopped, but the device is still physically connected then it will try to reset the port. | High | Application | Failed Testing | PROVU_V3_1_1_APP_0_4_2 PROVU_V_3.3.3_APP_v.0.6.3 3.4.2ADB |  | <WW 21Jul2026>: We tested this at D.G using ESD gun on version v0.6.3. Applying ESD strike to broncho camera. We see the broncho stream briefly stop, then recovers automatically. <WW 15Sep2026>: China reporting ESD strikes (air discharge) onto the bronchoscope shaft causes camera corruption. Refer to demo video: Demo_Video/v0.7.2 issues/broncho ESD failures 15Sep2026.mp4 |  |  | No
206 | closed | Updated VL/VS camera image processing settings are required | We need to update the default image settings for the VL/VS camera as follows: NOTE : These were defined using Ayush special version circa 22Jan2026. VL/VS : Gamma : 75 Brightness: -6 Contrast : 0 Saturation: -20 Sharpness: 0 Hue: 0 Colour Temperature: 0 Red: 0 Green: Disable 0 Blue: 0 Application image processing Pipeline priorities: colour convert: 1 horizontal flip: Disabled black level: disabled Gamma: 4 DeNoise: Disabled CLAHE: Disabled Brightness/Contrast: 2 Saturation: 3 Sharpness: disable Hue: Disabled Colour Temperature: Disabled Red: Disabled Green: Disabled Blue: Disabled | High | Application | Resolved |  |  | Davis <29-Jan-26>: Sharpness is part of the general control, and will not be disabled. Instead Sharpness will be set to priority 5 and value will be set to 0, which will not have any effect. Ayush <21-04-26>: The Gamma,hue, color … advance control shall be saved in admin profile with the priority. it shall not affect the standard user. Davis<23-Apr-26> Taken up in SRS changes Req ID #4, in Miletsone 4 <WW 16June2026>: These settings have been updated since and are now managed by the SRS content. | App: 0.5.0 |  | NA
207 | closed | File Delete not Working on single view | After Deleteing file from single view it switch to the next file then if I press delete button again the new file is not being deleted. | High | Application | Resolved | PROVU_V3_1_2_APP_0_4_3 PROVU_V_3.2.9_APP_v.0.5.9 | davis | Davis<23-Apr-26> Shall update in 4/29 release <WW 01Jul2026>: Tested on version 'PROVU_V_3.2.9_APP_v.0.5.9', see demo video '..v0.5.9 issues/files not being deleted-1.mp4'. If I enter a media folder and click on an image or video, then delete, it deletes this first one and the next image or video is shown.Then if I keep deleting the images or videos, it get's 'stuck' and does not delete the presently shown image/video and also does not let the user go forward or backwards to view other files. I have to go back out of the gallery and then back in again to make it work. <WW 17Aug2026>: Retested on v0.7.0 and fixed. | App: 0.6.9 |  | NA | Davis to check on this | 2026-08-14
208 | closed | Device get Reset during file view | When we tap on view on and View off for mutiple file in gallery page, the device hangs-Up and get reset . | High | Application | Resolved | PROVU_V3_1_2_APP_0_4_3 | davis | Davis<11062026> Need more details , couldnot reproduce the bug <WW 01Jul2026>: Tested using verison 'PROVU_V_3.2.9_APP_v.0.5.9' and could not reproduce this bug |  |  | NA
209 | closed | Bronchoscope Connection Failure | When using hardware version 0.1.5 with software version 0.4.0 we occasionally observe a white screen at startup after inserting the bronchoscope camera, followed by a connection error. We tested the same hardware with software version 0.3.6, and this issue was not observed. However, we tried the recent release version as well and observed the same error | High | Application | Resolved | PROVU_V3_1_2_APP_0_4_3 |  | Ayush<22-Jan-2026>: In current version 0.4.2, we are ensuring the camera is connected by check framerate >0, this is making detection too fast. Application configure the ov428 parameters before FPGA finish the configuration. if(.... && portAStatus->fps() > 0) ...) I recommend, we should check for fps()>25, this will ensure the camera is configured properly. However, this is not a standard way but we have to do this way as we don't have a handshaking with FPGA. <WW 01Jul2026>: Cannot see this specific issue on version 'PROVU_V_3.2.9_APP_v.0.5.9' | App: 0.4.4 |  | NA
210 | closed | Device Resets and boot the second partitions | Some time when device fails to boot or stuck in booting , so after restarting it loads the second partition and application | High | Yocto | Resolved | PROVU_V3_1_2_APP_0_4_3 | poornesh.g;#21 | Poornesh<11062026> Implemented in Yocto v3.2.5 | Yocto v3.2.5 |  | NA | Closed
211 | closed | Sleep Mode, background image visible | in sleep mode, I can still see the background image of the GUI. If there are no cameras attached, I can see a very feint image | High | Multiple | Resolved | PROVU_V3_1_2_APP_0_4_3 | davis | Davis<23-Apr-26> Updated in v4.9 App <WW 01Jul2026: Tested in version 'PROVU_V_3.2.9_APP_v.0.5.9' and resolved | APP_0.4.9 |  | NA
212 | closed | USB detection failure | something strange going on with the USB stick reading. I have a USB stick that I have been using for a while with no problem, then I use it in a display here and it is not recognised now. tried using a different USB stick and it is recognised. <WW 01Jul2026 post edit>: I have a good known working USB stick (USB3.0), attach to display, copy some media files to it, remove it and attach to a laptop and the laptop notifies me 'there is a problem with the USB drive, scan now to fix' type message. This only happens when I connect to 15" display and copy media to it. | High | Yocto | Resolved | PROVU_V_3_1_4_APP_0_4_5 PROVU_V_3.2.9_APP_v.0.5.9 | poornesh.g;#21 | 16-jun-2026:Updated USB rules <WW 01Jul2026>: See notes added today. I get issue on same USB stick after connecting to display and copying media files - then insert into laptop - Windows reports an error with the USB drive and recommends to 'scan to fix'. <WW 16Jul2026>: Re-tested on v0.6.1 and v0.6.2 and seems to be resolved. I can use USB stick for software upgrades and media copy and I do not get the 'USB stick corrupt' notification when I plug into a laptop. |  |  | NA
213 | closed | Low Disk Space Error with continuous audio | we have observed the low disk space warning with continous audio playing. | High | Application | Resolved | PROVU_V_3_1_4_APP_0_4_5 | davis.packiaraj;#18 | Davis<23-Apr-26> Updated in v0.4.6 App <WW 01Jul2026>: I have not seen this in recent versions. Tested version 'PROVU_V_3.2.9_APP_v.0.5.9' | APP_0.4.6 |  | NA
214 | CurrentPlan | Process Error | I did boot one device direct from power and seen "Process Error Media Capture Process is Not Running". I tried using screen recording but it is not fucntioning. I power cycled the device and the error not coming again I have still seen this on more revent software (app_v0.6.3). I simply powered up the device (from power off) initially with no cameras attached, then attached a VL camera, then Bronco camera. | High | Application | Ready for testing | PROVU_V_3_1_4_APP_0_4_5 PROVU_V_3.3.3_APP_v.0.6.3 | davis, poornesh | Davis<23-Apr-26> Need to reproduce the Bug and check, need Ayush help here. Ayush: Test with the same version to reproduce the issue <Davis 22 JUl 2026> : Trying to replicate the bug <Devi 24 JUL 2026> : Unable to replicate the bug <Poornesh/Devi 12 Aug 2026> Launcher script to be updated by chekcing the status of the processes post execution . | App:v0.7.2 |  | Yes | Check the code | 2026-08-18 | 2026-08-26
215 | closed | Reboot bootloader command failing | I tried the device to switch to load the wic image but reboot bootloader command not working | High | Application | Resolved | PROVU_V_3_1_4_APP_0_4_5 |  |  |  |  | NA
216 | closed | Unexpected workflow during software upgrade | I have a valid .rauc file on the USB stick. Attach to device and click on software upgrade. Click on 'Download software', but then after it successfully downloads, leave the device to timeout. The device will return to the home screen (no alerts given). This is ok, but if I remove the USB stick, re-attach, then go through the software upgrade process again, it does not appear to download the software from the USB, but proceeds straight to installing the previously downloaded software. This is not intuitive. If the user re-starts the software upgrade process, it should always download the file from the USB stick first, before installing. | Low | Application | Resolved | PROVU_V3.1.3-ADB-broncho-i2c4-test PROVU_3.3.2_APP_V_0.6.2 | davis.packiaraj;#18 | Davis<23-Apr-26> Looking in to the issue, will update ASAP <WW 22Sep2026>: Tested in v0.7.7 and resolved |  |  | No | Davis to check on this | 2026-08-18 | 2026-08-26
217 | closed | Software Upgrade | After flashing the new software and restarting the device, it works correctly. However, if power off and restart after 1 minute, the software will revert back to the old version. It requires flashing 2~3 times to become stable. | Low | Yocto | Resolved |  |  | 28-April-2026: These issue reported by norton during test in DG 16-Jun-2026:Could not able to replicate in our device.Requsting FC team to check this in latest release <WW 01Jul2026: Could not replicate this issue in latest software tested 'PROVU_V_3.2.9_APP_v.0.5.9' |  |  | NA
218 | closed | FAN Control | The cooling fan does not operate | Low | Application | Resolved | PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 | 28-April-2026: These issue reported by norton during test in DG Davis<11062026> Looking into the issue Davis<19062026>: Fan script removed from the init. Application will take control of fan. <WW 01Jul2026>: Tested on version 'PROVU_V_3.2.9_APP_v.0.5.9'. The fan pulses on briefly during boot-up, but does not seem to operate when battery is being charged, or discharge even for elevated internal temperature; fan is OFF. this is not meeting present requirements. Davis<6-July2026> We shall have test S/W with temp range changed to test the functionality . Davis<14072026> Implementation done in V0.6.2 , testing in progress to be released today <WW 16Jul2026>: Confirmed the fan does now operate. I have not tested the specific requirements of temperature<>fan-speed - this will be done in first iteration of testing. | App v 0.6.2 |  | Yes
219 | CurrentPlan | Display flickering | After long-time operation, when temperature rises, partial areas of the screen background display flicker | Low | Hardware | In progress |  | China | 28-April-2026: These issue reported by norton during test in DG |  |  | Yes | NA | NA | NA
220 | closed | Camera Detection | UI software issue: Sometimes the USB camera displays image directly after connection, but sometimes it is necessary to press the virtual B/C camera button on the screen to activate and display the connected camera image | Low | Application | Resolved | PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | 28-April-2026: These issue reported by norton during test in DG Davis<11062026> Looking into the issue Davis<25062026> Need some clarification on the bug--Shall discuss with Ayush and come back Davis<30062026>: Change in requirement , requested a write-up from FC team. Davis<20720206> Details received , WIP Davis<6-July2026> To be released in v0.6.0 on 07/7 Davis<9-July-2026> To be resolved in next v0.6.1 release on 7/14 Davis<14072026> Implementation in progress trying to update in v0.6.2 done to be released today Davis<16072026> Have some doubt on the details shared by Abhishek , to be discussed in today's 2pm call and get clarified. Davis<17072026> Based on the new requirement, implementation done <WW 20Jul2026>: Tested this again on v0.6.3 and it is still not consistenyl showing dual-view by default when two cameras are attached. This is not easy to reproduce, but by cycling through different combinations of inserting/removing cameras, I can see this fault. <Davis 23 JUl 2026> : Checking for Root cause <WW 24Jul2026>: Tested in v0.6.4 and resolved |  | Need to confirm and updated GUI spec will be shared | Yes
221 | closed | Software upgrade - boot screen freezing | Sometimes after I perform a software upgrade using USB, the reboot sequence happens, but it freezes on the boot-up splash screen. As an example, I already had software V3.1.9 installed and was updating to V3.1.9 (not a typo - I was intentionally trying to upgrade to the same version!). It successfully downloaded and installed the software, provided the prompt for reboot after 'success', but when it rebooted, it got 'stuck' on the boot-up splash screen. To fix, I had to press+hold the power button which provided a reset, it then booted up correctly. | High | Multiple | Resolved | PROVU_V_3_1_9_APP_0_4_9 | poornesh.g;#21 | 16-jun-2026:Requsting FC team to check this in latest release <WW 01Jul2026>: Tested on version 'PROVU_V_3.2.9_APP_v.0.5.9' and cannot reproduce. |  |  | NA
222 | closed | Camera image stuttering on startup | Sometimes, when I start the display from 'cold' and use a camera, the image is stuttering. There is some sort of intermitent lag, but the frequency of the lag seems to be constant. | High | Hardware | Resolved | PROVU_V_3_2_1_APP_0_5_1 | China | Davis<11062026>Looking into the issue WW <11June2026> : Ayush reported this is an FPGA issue. China is looking into it. <WW 23Jul2026>: China released new FPGA version v.9 that so far is resolved this issue. FPGA version has been uploaded in Azure. |  |  | Yes
223 | closed | Possible USB stick becomes corrupt after removal from display | We have had instances where a USB stick has become corrupt after using it and removing it from the display. To fix, we had to re-format the USB stick. It is curious that some people are also noticing that on removal of the stick from the display, when they plug into a laptop, it says the USB stick has a problem and should be checked.. | High | Application | Resolved | PROVU_V_3_1_9_APP_0_4_9 | davis.packiaraj;#18/Poornesh | Davis<11062026> Need to discuss internally with Yocto team and update Davis<19062026>: Poornesh to modify the usbrules file, so that application can mount/unmount the usb when using it. Davis<25062026>: Estimated closure on 7th July-tuesday, Yocto to remove auto mount feature from build | App: v0.6.0 |  | Yes
224 | closed | Software upgrade GUI and workflow is broken | On latest software, if we try to update the software with another version, it scans the USB stick, then gets to download page and freezes, then reports an error with the update. But then if you return back to the software update screen (without re-starting and without removing the USB stick), it successfully downloads and installs the software. | High | Application | Resolved | PROVU_V_3_2_1_APP_0_5_1 | davis.packiaraj;#18 | Davis<11062026>Looking into the issue Davis<25062026> We need to modify the S/W upgrade flow to use CPP instead of QML- Closure by June 30th 2026 Davis<6-July2026> To be released in v0.6.0 on 07/7 <WW 16Jul2026>: Tested on v0.6.2 and this issue is fixed | App v 0.6.0 |  | Yes
225 | closed | Display flickering | Upon multiple power cycle the display start flickering. It only get stablised onece device powered down and wait for 10 seconds and power up | High | Hardware | Resolved | PROVU_V_3_2_1_APP_0_5_1 |  | Same issue as 222 above. |  |  | NA
226 | closed | Broncho auto-gain adjustment speed | The auto-gain feature on the bronchoscope is too slow. We have had feedback from clinicans on recent software that this needs to be fast because the clinicians tend to move the scope quite quickly in the anatomy. I believe this is an OV428 setting / camera-chip settings that can be adjusted - but we need to be careful about oscillations. | High | Application | Resolved | OS Build Version V3.19-USB Application Version 0.4.9 | Wayne + Ayush | Davis<11062026> @Ayush , Will the new OV428 settings shared , resolve the issue? <22062026>: New settings updated <22062026>: WW - These settings have been updated again and need to be implemented in software ('Final1') Davis<25062026>: Wayne can you please share the new settings . once shared shall update the S/W in next release Davis<30062026>: New settings received from Wayne on 6/26 , applied and released in v5.9 <WW 01Jul2026>: Tested on version 'PROVU_V_3.2.9_APP_v.0.5.9' and seems to include new settings | App:v0.5.8 |  | Yes
227 | closed | Broncho emergency mode is still seen on screen in sleep mode | Feedback from demo unit in USA - if the user has broncho attached and puts the device into sleep mode, it proceeds into sleep mode, but the broncho emergency mode image is still shown..? | High | Application | Resolved | OS Build Version V3.19-USB Application Version 0.4.9 | Davis | Davis <11062026> Need to discuss with FC team |  |  | NA
228 | closed | Sleep mode, image corruption when cameras remain inserted | If I insert cameras (broncho or USB, or both), then go into sleep mode, it enters sleep mode, but when I wake it up again, there is image corruption and multiple other side-effects: 1. If I subsequently remove the cameras, the last image remains frozen on the display and the static image is shown overlaid onto the frozen image. 2. For broncho, I sometimes see corrupt pixels in the right-hand side of the display | High | Application | Resolved | PROVU_V_3_2_1_APP_0_5_1 PROVU_V_3.2.9_APP_v.0.5.9 |  | Davis<11062026> Resolved the issue w.r.t USB camera in App v0.5.5 <WW 01Jul2026>: still getting corruption (but not all the time) when I wake from sleep mode with cameras attached. Can be broncho or USB cameras. <WW 24Jul2026>: this is resolved in v0.6.4. | App:0.5.5 |  | NA
229 | closed | Some events are not being logged in the event log | feedback from USA demo : A user event was generated when 3 cameras were attached; the software correctly displayed the pop-up warning to only use 2 cameras. This was ok, but this event was not recorded in the event log | Low | Application | Resolved | PROVU_V_3_1_9_APP_0_4_9 PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 | Davis<11062026> Updated in App v 0.5.4 <WW 01Jul2026> : Tested on version 'PROVU_V_3.2.9_APP_v.0.5.9' and this event is not getting logged in the event log. <WW 17Aug2026>: Tested in v0.7.0 and fixed | App v 0.5.4 |  | NA | Davis to check on this | 2026-08-11 | 2026-08-14
230 | closed | auto switch into dual-view mode when 2nd camera attached | When there is only one camera attached and just one view is shown. If the user attaches a second camera, the device should automatically switch to dual-view mode. At the moment, the software correctly recognises the second camera, but keeps single view. When I have broncho inserted and viewing image, then insert VL camera in PortB, it does not automatically go into dual-view mode. It recognises the insertion of VL camera in PortB but beeps and keeps in single view but shows the icon on PortB 'ready to be used'. BUT - on two USB cameras (VL and VS), this seems to work sometimes (not all the time!). Its just if I have broncho as one camera, then insert another camera into portB, or portC, this auto switch to dual-view does not work as expected | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 | davis.packiaraj;#18 | Davis<19062026>: This issue is related to #220. Will look into it. Davis<25062026> Duplicate bug, already captured against SL#220. <WW 01Jul2026> : closed due to duplicate bug #220 |  | Need to confirm and updated GUI spec will be shared | Yes
231 | closed | false broncho voltage error | After power-up, with broncho camera attached, I get 'Broncho Supply Error : broncho isolated power supply voltage is out of range (3.3V …). Logged in as admin for this | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 | davis.packiaraj;#18, Poornesh | Davis<19062026>: Currently measuring 1.65 V , as per req 3.3 =/- 12%. <22062026>: Threshold changed. Known Issue: Once Broncho is disconnected and connected back, we do not have access to adc from software. Poornesh to look into it Davis<25062026>: Yacto shared the sys commands, to be added in next release on 6/26 <WW 01Jul2026>: Tested in version 'PROVU_V_3.2.9_APP_v.0.5.9' and do not see this error | App: v0.5.9 |  | Yes
232 | closed | portA connection icon incorrect colour | with no cameras attached (I just removed the broncho), the portA icon is different to the others - the crossed out camera icon is not highlighted in grey but the others are?) Logged in as admin for this | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 | davis.packiaraj;#18 | Davis<19062026>: Working on the issue <WW 01Jul2026>: Tested on version 'PROVU_V_3.2.9_APP_v.0.5.9' and resolved | App:v0.5.8 |  | Yes
233 | closed | broncho image corrupted or not shown | I plugged in a bronchoscope and the image did not come on the screen initially, then it got corrupted image - it said 'Camera Error Replace ProVu input device' Logged in as admin for this | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 | davis.packiaraj;#18 | Davis<19062026>: Need Ayush help Davis<02072026>: Added Camera re-connection during camera error <WW 21Jul2026>: Tested software v0.6.3 on multiple devices (using v1.6 PCB) and so far there are no corruption issues with broncho during insertion. | App:v0.6.0 |  | Yes
234 | closed | broncho image quality | image quality of broncho is not good enough - glare - too much auto-gain/exposure etc. Logged in as admin for this | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 | Ayush | <WW 01Jul2026> : Confirmed fixed in version "Yocto_v3.2.9/App_v0.5.9". Contains updated OV428 register values provided by Ayush. |  |  | Yes
235 | closed | low disk space error | I got a 'low disk space' error when I attached a VL camera. I had already removed the broncho. To note, I had used the 'ADB' version and had re-flashed the entire image using upgrade_tool for this test. Logged in as admin for this | High | Yocto | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 PROVU_V_3.3.1_APP_v.0.6.1 | Poornesh | Poornesh<25062026> Couldn't reproduce the error as of now. Poornesh<2072026> As Wayne updated revised binaries , this error is not occuring. <WW : 13Jul2026>: List of commands received from Poornesh on email (10th July2026) helped to resolve the issue (with one adjustment of sequence of commands). Poornesh will bake this solution into the Yocto layer boot-up to prevent this happening in future software builds. <Poornesh/Devi 22 JUl 2026> : Trying to replicate the bug, working out options to resolve, post v6.4 will be released <WW 24Jul2026>: Still seeing this in v0.6.4. I flashed the adb (.wic) version (previous software was v0.6.3-ADB), and on power-up I am seeing 'Low disk space'. <Poornesh/Devi 12 Aug 2026> Implemented in v6.6 , FC to test and come back <WW 22Sep2026>: Tested in v0.7.7 and resolved | Appv 0.6.6 |  | Yes | ok-open-Ayush to update SRS | NA | NA
236 | closed | bronchoscope image corruption when re-inserted | I re-inserted broncho and had image corruption issu not get it to correct itself by removing/re-inserting multiple times. I had to reset the device. Upon reset, it did display the broncho image correctly (I had broncho attached from startup), but now I get error 'incompatible bronchoscope' shown - saying that it was not S,M,L version. Then the 'broncho supply error : ..' error showed and dominated the error window. I could not see the previous error..? I thought if we had multiple errors at same time, then user should be able to see these scrolling through on a timer? - post-edit - sometimes it does cycle through the various errors that have been detected, but not all the time. Logged in as admin for this | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 | Wayne / Ayush | <WW 16Jul2026>: Closing this out : tested on v0.6.2 and so far are not seeing corruption on bronchoscope when repeated cycling through removing/re-inserting camera. | App: v0.5.8 |  | Yes
237 | closed | image adjustment icon colour | Logged in only as guest account - when I have broncho only attached, then adjust image settings, when I 'save and close', the image settings are applied correctly, but the icon at the top border (showing the image adjustment button) is still grey - this needs to be blue if the settings have been changed - even for Guest. For guest account, the settings must be able to be changed, but not saved to the guest account after the device has been removed, or device is powered off then back on. | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 | davis.packiaraj;#18 | Davis<19062026>: Working on the issue Davis<25062026>: To be updated in Tuesday App release v0.6.0, 6/30 Davis<6-July2026>To be released in v0.6.0 on 07/7 Davis<14072026> Details received from Abhi, implementation done in V0.6.2 to be released today <WW 20Jul2026>: Resolved in v0.6.3 | Appv 0.6.2 |  | Yes
238 | closed | false LED current check error reported after power up | When I powered up with broncho already attached, I got a error 'LED current low : LED current draw voltage is below the minimum threshold (2.2V)'. | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 | davis.packiaraj;#18, Poornesh | Davis<19062026>: Working on the issue. <22062026>: Threshold changed. Known Issue: Once Broncho is disconnected and connected back, we do not have access to adc from software. Poornesh to look into it Davis<25062026> To be updated in 6/26 release v0.5.9 Davis<30062026> Released in v5.9 <WW 01Jul2026>: Tested in version 'PROVU_V_3.2.9_APP_v.0.5.9' and do not see this msg | v0.5.9 |  | Yes
239 | closed | camera brightness control adjust graphic issue | when I adjustthe brightness to max with broncho attached (moving slider to right), I get a strange blue bar continuing beyond the right of the brightness control. | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 |  | <WW 16Jul2026>: Tested version v0.6.2 and this has been resolved. |  |  | No
240 | closed | guest accounts is saving image settings | logged in as guest, broncho attached, then change the brightness - all ok. Then I remove the broncho, wait a few seconds - comfirmed the device has recognised the broncho has been removed, then I re-insert the broncho and check the image and settings - the brightness is still set to the changed value - for guest accounts, the image adjustments must be set back to default if the broncho or other cameras are removed. | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 |  | <WW 16Jul2026>: Requirements have changed and we now have this 1min dwell period after a camera has been removed to retain some configuration for that port (viewport location, zoom, streaming status, image controls for Guest). This seems to be ok on latest software v0.6.2 |  |  | No
241 | closed | image adjustment icon colours - 2 | When I adjust the image settings the icon colour of the image adjustment control in the top blue border is not behaving correctly. When you go into the adjustment settings menu, it should go dark blue, then stay dark blue when you are in the menu, then if you save+close, if the settings have changed, the icon at the top shoud go light blue. Presently, it is going dark blue on entry to the menu, then changes to light blue (when the menu is still visible) if a setting is changed, then goes back to grey when you save+close. this was the same for all camera ports. | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | Davis<19062026>Looking into the issue Davis<25062026>To be updated in Tuesday App release v0.6.0, 6/30 Davis<02072026> Need to discuss in today's call regarding req details conflict in GTD version C and SRS version B As per discussion with Wayne on JUly 2nd Bug put on hold, Abhishek to confirm Davis<14072026> Details received from Abhi, implementation done in V0.6.2 to be released today <WW 22Jul2026>: On v0.6.3, when you enter the image adjustment menu, the icon turns dark blue - correct. But when you change some settings, the icon changes to light blue when still in the menu. this is not correct. it should stay dark blue. If you exit the menu and the settings have changed, the icon should go light blue - this is now correct. If you do not revert to defaults and exit the menu, then it should go back to grey - this is also correct. See Demo Video '.../Demo Videos/v0.6.3 issues/bug241 demo.mp4' <Davis 22 JUl 2026> : Working on the bug , to be relsoved in v6.4 today's release. This change to be updated in SRS. <WW 24Jul2026>: Tested on v0.6.4 and resolved. SRS to be updated | App v 0.6.4 |  | Yes
242 | closed | audio feedback needed when changing tones using the slider | When we adjust audio settings, there is no feedback as it gets adjusted - this should be added at least for the keytones. Suggest playing the audio tone when the slider control is released each time? | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 | davis.packiaraj;#18 | <WW 06Aug2026>: Some points: 1. Please remove this tone from all other sliders - only use on the 'Tone Volume' slider. (e..g I now get a tone when adjusting image control sliders, etc, please remove). 2. On the 'Tone Volume' slider - if I set the volume to '0', it correctly turns off the touch tone - good, but now when I increase the Tone volume slider, the first press/release I do not hear the tone. <WW 17Aug2026>: retested om v0.7.0 and fixed | App v 0.6.9 |  | No | Davis to check on this | 2026-08-11 | 2026-08-14
243 | closed | alert and tone volume sliders | Alert volume' and 'Tone volume' are the wrong way around in settings | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 |  | <WW 01Jul2026>: Closed due to repeat of bug#21 |  |  | No
244 | CurrentPlan | audio tone minimum setting too low | The tone volume only seems to be heard when the minimum level is about 6.? We can't heard it much when the setting is below this. We do need a way to disable the tones, so suggest using level 6 as minimum, but have one setting below this which disables the tone (or sets volume to 0). | High | Application | Ready for testing | ProVu v3.2.6-ADB / App V.0.5.6 | davis.packiaraj;#18 |  | App v 0.6.9 |  | No | Davis to check on this-Alert tone 6 to 10 and tone volume 0, 6 to 10 | 2026-08-11 | 2026-08-14
245 | closed | video shows blank when scrolling through videos in the media gallery | in media gallery, when we scrolling through subsequent videos, the 'PLAY' word stays shown, even when the videos briefly go blank as the software is loading the next video into the media player. | High | Application | No fix required | ProVu v3.2.6-ADB / App V.0.5.6 | davis.packiaraj;#18 |  |  |  | No | Use as it is- No change in SW
246 | CurrentPlan | error pop-up window has overlapping text | on the bottom fault pop-up, some text which is long is overalapping the 'X' button on the RHS (of the fault pop-up window) | High | Application | Ready for testing | ProVu v3.2.6-ADB / App V.0.5.6 | Ayush |  |  |  | No | Test latest version and close | NA
247 | closed | incorrectly shows dual-view mode when I close one camera | When I have three cameras attached, then close the left viewport camera, the third camera is automatically shown. scenario is this: I remove all cameras, then insert camera to PortC, confirm single view, then insert camera to PortB, confirm dual-view, then insert broncho to PortA - confirm there is no change to the dual-view. Then I use the GUI to close the left-viewport (PortC) - what happens then is the broncho is automatically shown in the left-viewport - which is not correct. In this case, if I close the PortC view, it should go to single viewport (PortB), and hold PortA and PortC waiting for user to select. | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 | davis.packiaraj;#18 | Davis<19062026>: Working on the issue Davis<25062026>: Related bug to SL#230 & 220 <WW 22Jul2026>: Tested in v0.6.3 and resolved. | App: v0.6.3 | Flexi will provide scenrio examples to davis | Yes
248 | closed | zoom is incorrectly filling the screen (changing viewport size) | If I use the USB camera (portB or PortC), and use the zoom-in button (at the top of the blue border), the image fills the viewport - this is not right - it should zoom but keep the viewport size. This is different behaviour to the'max view' button (bottom left/right) | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 | davis.packiaraj;#18 | Davis<19062026>: Working on the issue Davis<25062026> Putting this bug on hold unless Wayne confirms on signle view port for dual camera . This feature will completely change if proceeding with single view port. If not going with single view port shall update on 6/30 Appv0.6.0 release. else update with single view port release which will take 10days from Wayne's confirmation date. (July 10th) <WW 22 Jul2026> : This has been tested on v0.6.3 and now resolved. | App: v0.6.0 |  | Yes
249 | closed | conflict between max-size button and zoom button functions | I use dual-cam PortC + PortB (C is on left), use normal viewport size and no zoom (so the max size buttons on the bottom left/right are grey and the image is native resolution and not filling each view port. All good. Now I select the max-size button for one of the ports (bottom right hand button). This button then changes blue colour - all good. Then I close the left viewport (pressing the top 'X' button on top border) and the single image is shown for PortB. The image shown in single view is not right - it is already zoomed in, and there are not icons showing that it is zoomed in. Now I zoom out using top border control and now it shows correct image size. But now I re-enable PortC (using icons on the left port) - and the right viewport (PortB) now has the max-size button coloured in Blue. | Low | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 | davis.packiaraj;#18 | Davis<19062026>: Working on the issue Davis<25062026>: Putting this bug on hold unless Wayne confirms on single view port for dual camera . This feature will completely change if proceeding with single view port. If not going with single view port shall update on 6/30 Appv0.6.0 release. else update with single view port release which will take 10days (July 10th) from Wayne's confirmation date <WW 22Jul2026>: Tested in v0.6.3 and resolved. | App: v0.6.0 |  | Yes
250 | closed | sleep mode issues - 2 broncho image is briefly displayed in emergency mode (FPGA) | I have broncho attached and working normal. I put the device into sleep mode and it does transition into sleep, but the 'emergency FPGA' image is shown briefly before switching off. Is it possible to switch this broncho interface port off first, then switch the LCD off? | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 |  | <WW 16 Jul2026> : Fixed in v0.6.2 |  |  | No
251 | closed | broncho image is corrupt on wake-up from sleep | When waking up from sleep mode, the broncho camera is sometimes being corrupted. Most times, if I have bronchoscope already attached and I go into sleep mode, then wake up, the image freezes. I have to remove the broncho and re-insert to correct. Sometimes, when I re-insert the broncho the image is corrupted and cannot recover. To recover, I need to power-cycle the device. <22 June 2026 - WW> : Latest software (v0.5.8) is better than v0.5.6 but still has issues coming out of sleep mode. Two main issues: 1. If I have broncho + USB camera attached, then put into sleep mode, then during sleep mode remove the USB camera, when it comes back out of sleep mode, the broncho camera is corrupted and switched to emergency FPGA mode. 2. When I have two cameras (USB + broncho) already attached, then put into sleep mode, I keep the two cameras attached, then bring it out of sleep mode, I get a very brief flash of the previous dual camera screen with 'Going to sleep' pop-up shown. I have taken some video of the above here : https://flexicaremedicallimited.sharepoint.com/:f:/r/sites/ProVu2MindteckShare/Shared%20Documents/General/Demo_Video/v0.5.8%20issues?csf=1&web=1&e=BT3y6L | High | Application | Resolved | ProVu v3.2.6-ADB / App V.0.5.6 ProVu v3.2.8-ADB / App v0.5.8 PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 | Davis<26062026> :To be releasedon 6/26 S/W App v0.5.9 Davis<30062026> Released in App v 0.5.9 <WW 01Jul2026>: latest version 'PROVU_V_3.2.9_APP_v.0.5.9' still has camera image corruption when waking from sleep Davis<6-July2026> Resolved, shall be released in v0.6.0 App on 07/7 <WW 21Jul2026>: Retested on software v0.6.3 (and hardware V1.6) and so far have not seen broncho corruption from power-up or sleep mode. | App v0.6.0 |  | Yes
252 | Backlog | Gallery view full-screen view confusion with live view | when on gallery view and viewing snapshot images, then press full-screen, these could be confused with live images - need some way to show these are recorded. Flexicare to design these | Low | Application | Ready for testing | ProVu v3.2.7-ADB / App V.0.5.7 | davis.packiaraj;#18 |  | App: v0.8.0 |  | No | Inputs from Flexicare awaited | Inputs from Flexicare awaited | NA
253 | closed | cannot select next/previous media files when in full-screen mode | on full screen mode when viewing snapshots, you can't go forward/back using any buttons - no buttons available. You have to go back to normal size view | Low | Application | No fix required | ProVu v3.2.7-ADB / App V.0.5.7 |  |  |  |  | No | NA
254 | CurrentPlan | settings icon colour when in gallery menu | when in gallery view and I press the settings button on the RHS, it highlights the settings button and takes a long time to switch into the menu.. | Low | Application | No fix required | ProVu v3.2.7-ADB / App V.0.5.7 | davis.packiaraj;#18 | 26Aug2026: Loading Meadia Veiwer/ Gallery components takes time, which is expected. No fix required |  |  | No | Davis to check on this | 2026-08-18 | NA
255 | CurrentPlan | unexpected action when touching pop-up during software completion confirmation | when in gallery mode and I have selected files to copy to USB, then I proceed to copy the files and all is ok. After it completes, I get the completion confirm pop-up 'File transfer complete'. When in this pop-up, if I touch some areas of the pop-up (NOT the close button), then one of the video or snapshot files seems to get loaded in the background main window in the gallery. I can see the filename appear in the top of the screen above the pop-up box. <WW post edit 24Jul2026>: Also, if you are copying a large file(s) that takes time, when the copy pop-up box is visible, you can press the other buttons in the background. You can nagivate back to the home screen and you can press 'DeleteAll' buttons. This is not intuitive and will cause problems | Low | Application | Ready for testing | ProVu v3.2.7-ADB / App V.0.5.7 | davis.packiaraj;#18 | <WW 17Aug2026>: tested on v0.7.0. When the pop-up is in shown, I can still press buttons in the background which are being actioned (e.g.delete All still works). <WW 22Sep2026>: Tested v0.7.7 and failed. Some of the buttons in the background are disabled, also the power button also seems to be disabled, but I can still press the 'home' button when the pop-up is active and it moves back to the home screen. | App v 0.8.2 |  | No | Davis to check on this | 2026-08-14 | 2026-08-17
256 | closed | Unexpected behaviour during software upgrade process with no USB stick attached | software upgrade screen - if there is no USB stick attached, I can press the button to start software update and it starts scanning - but there is no USB stick attached - so why does it need to start scanning the stick that is not yet attached? | Low | Application | Resolved | ProVu v3.2.7-ADB / App V.0.5.7 | davis.packiaraj;#18 | <WW 17Aug2026>: tested on v0.7.0 and fixed - if there is no USB stick attached, I see the pop-up 'No compatible USB device found' | App: 0.7.0 |  | No | Re-test with latest version-Davis to check and update | 2026-08-14
257 | Backlog | software upgrade failure screen | 1. if there is a failure to do software update, the failure screen needs work done. Are we sure about the colour and size of the font? seems dramatic?. 2. if there has been no USB stick attached and it has failed due to no USB stick attached, or no software downloaded, why does it have to reset the whole display? | Low | Multiple | Require inputs from Flexi | ProVu v3.2.7-ADB / App V.0.5.7 | Abhishek |  |  |  | No | Update GTD with screen spec, Davis to implement the same. No new screen at this stage | NA | NA
258 | CurrentPlan | power button still works during USB file transfer | when transferring media files (or other files) to USB stick, we should ignore the power button pressed or tapped - presently it will go into sleep mode during copying files to USB stick and then we don't know if it copied all the files. | Low | Application | Ready for testing | ProVu v3.2.7-ADB / App V.0.5.7 | davis.packiaraj;#18 | <WW 22Sep2026>: Tested in version v0.7.7. Partly resolved. The power button is disabled during the USB transfer, but when the final pop-up is still showing ('File transfer complete'), I can still press power button and other buttons in the background. The 'requirement' needs to be that any window that has a higher 'z' order and is in focus needs to be cleared before the user can press the other buttons in the background. | App v 0.8.2 |  | No | Davis to check on this | 2026-08-18 | 2026-08-26
259 | closed | power button still works during software updates | we should disable the power button during software download/install. Presently I can put the device into sleep mode when software is downloading - we should prevent this. | Low | Application | Resolved | ProVu v3.2.7-ADB / App V.0.5.7 | davis.packiaraj;#18 |  | App v0.7.2 |  | No | Davis to check on this | 2026-08-18 | 2026-08-26
260 | closed | hardware version number displayed is not as expected | When I look at the hardware version number shown in the information GUI page, it is not as expected. I see a 3-digit number, but the expectation should be a format like 'A.1', or 'B.3'. This table should be in the SRS to define the various digital and analog values attributed to each version letter/number | Low | Application | Resolved | ProVu v3.2.7-ADB / App V.0.5.7 |  | Davis<26062026> : To be added in 6/26 release v0.5.9 <20 Jul2026>: Resolved in v0.6.3 |  |  | No
261 | closed | bronchoscope image default settings are incorrect | I just re-programmed the device using ADB and using the latest v0.5.8 version. I setup the admin account as usual. I reset the device after initial configuration. After reset, I attached a bronchoscope, then logged in as Admin. Then I pressed the 'image settings' button, and the image controls were set to the USB camera, not the broncho camera. e.g. Gamma=75, brightness=-6, contrast=0, saturation=-20, Red=12, Blue=20. BUT - I am still also getting the 'low disk space' error. Is there something strange going on with the partitions on this device that has corrupted the user account settings?? | High | Application | Resolved | ProVu v3.2.8-ADB / App V0.5.8 | davis.packiaraj;#18 | Davis<26062026>: Need to reproduce the bug , shall be corrected in S/W v0.6.0 release dated 6/30 Davis<30062026> Wayne shared the details on 6/26, implemented and released in App v 0.5.9 <WW 16Jul2026>: image control settings are on defaults from power-up unless Admin account has already adjusted them. - tested on v0.6.2 |  |  | Yes
262 | closed | media gallery search text does not get cleared on re-entry | If I set the media gallery search text to some search string (e.g. "2026-03-21"), if I exit the menu, then go back into the gallery, the same text remains in the search bar. | High | Multiple | Resolved | ProVu v3.2.8-ADB / App V0.5.8 | davis.packiaraj;#18 | Davis<30062026> Shall be released in v 0.6.0 dated 1/7 <WW 21Jul2026>: Resolved in version v0.6.3 | App: v0.6.3 |  | No
263 | closed | brief flash of USB port error when USB camera is attached | when I insert a USB camera, I get the notification tone immediately (very good), but there is a very brief flash of an error pop-up at the bottom of the screen. I cannot see what the error text is, but it is there. This should not be happening in normal case. | High | Multiple | Resolved | PROVU_V_3.2.9_APP_v.0.5.9 |  | <WW 16Jul2026>: Tested version v0.6.2 and this has been resolved. |  |  | No
264 | closed | Hardware version ID is fluctuating between two values | In Information page, I look at the HardwareID and it is displayed in the correct format (e.g. 'A.1'), but is fluctuating between 'A.1' and 'A.2' periodically. Video uploaded '../v0.5.9 issues/HardwareID read unstable.mp4' | High | Application | Resolved | PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 | <Poornesh/Devi 14Aug2026>: Issue got resolved Root cause: Delay is not proper b/w the ADC powerup and starting conversion. Delay has been set up for 16clk period instead of 8 clk period <WW 20Aug2026>: Tested v0.7.1 and fixed (tested on v1.6PCB) |  |  | No | Gopi to verify thru HW | 2026-08-18 | 2026-08-19
265 | closed | Software version shown during upgrade before installation | during the software upgrade, I can download software and the application now shows the version number of the upgrade package on the USB stick before I can proceed to install it. This is all good, but the format and version number is not what I was expecting. It seems to be using the Yocto version number, instead of the application version number. We need to decide which version number is used consistently, or show both. | High | Application | Resolved | PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 | <WW 22Sep2026>: Tested in v0.7.7 and resolved. We agreed to use the Yoto build version and software has implemented this. | App v 0.6.9 |  | No | Update GTD 989 & Davis to implement the same. | 2026-08-11 | 2026-08-14
262 | closed | bronchoscope incompatible version flash on screen during insertion | When I insert a bronchscope, I get a brief flash up of the error 'incompatible Bronchoscope', then it goes away. See demo video '../v0.5.9 issues/bronco imcompat error flash.mp4' | High | Application | Resolved | PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 | Davis<6-July2026> As per 7/02 call , Mindteck to mask the error, to release in v0.6.0 on 07/7 <WW 16Jul202>: Have not seen this in latest version v0.6.2 | App: v0.6.0 |  | Yes
263 | closed | order of the event log is wrong | The events in the event log are not ordered correctly. For some reason, I have a set of events that were logged on '2026-07-01' but the list does not seem to order these with the most recent time-stamp first in the list. There does not seem to be any order. It is difficuly to explain the result without listing all the event so pls refer to demo video : '../v0.5.9 issues/event list order.mp4'. It has resulted in any new events being put at the bottom of the list, even when the date and time are more recent than the existing entries in the list.. | High | Application | Resolved | PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 | <WW 22Sep2026>: tested in v0.7.7 and resolved. | App v0.7.2 |  | No | Introduce incremental ID's to the logs . To capture the last event index for ordering in seq. Use a circuilarQ type mechanism to store and show events. Most recent event needs to be shown at the top of the list when viewing in the GUI | 2026-08-18 | 2026-08-26
264 | closed | audio not working sometimes | Sometimes after boot-up, the audio is not working. Keytones not heard, removing/inserting cameras or PSU does not trigger an audio file. I have tried logging in as Admin, or normal guest user, but no audio. Note that I had HDMI monitor connected at this time. To fix, I need to power-cycle the display. sometimes it works. | High | Application | Resolved | PROVU_V_3.2.9_APP_v.0.5.9 | Yacto | Poornesh/Devi<6-July2026> To release in Yacto v 3.3.0 on 07/7 <WW 16Jul2026>: Tested on v0.6.2 and have not seen this issue again. |  |  | Yes
265 | closed | permits creating new accounts even with incomplete fields after timeout | I can start to create a new user accounts (logged in as Admin), I created a firstname, surname and username, but then leave some fields blank - e,g, password field is blank. Then do not touch the touchscreen (still has the account creation page visible) and wait for it to timeout (return to home screen after 2mins). The partially created account will be added to the list and it will leave the password fields blank. | High | Application | No fix required | PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 |  |  |  | No | NA
266 | closed | event logs are getting erased when I do a factory reset | Event logs are getting erased when I do a factory reset. Event logs and error logs should not get erased after a factor reset | High | Application | Resolved | PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 | <WW 07Aug2026>: Tested in v0.6.7 and fixed | App v 0.6.7 |  | No | Davis to check on this
267 | closed | broncho camera corrupted sometimes when attached | Sometimes when I connect the broncho camera, the image is corrupted or not shown. Refer to demo video '../v0.5.9 issues/broncho corrupt-1.mp4' | High | Application | Resolved | PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 | Davis<6-July2026> This error is happening irrespective of camera streaming from application. We have used G streamer to stream the camera and could able to reproduce the same issue. Hence this is pointing to HW concern Davis<7-July2026> adc/eeprom i2c binding is blocked on the driver layer until fpga is ready and configured in ov428 bus. <WW 21Jul2026>: Retested in v0.6.3 on multiple displays and no corruption so far. | App: v0.6.0 |  | Yes
268 | CurrentPlan | device does not detect missing touchscreen comms on startup | The software should detect a touchscreen device attached on startup. When I remove the touchscreen cable, there is no error displayed | High | Application | Ready for testing | PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 |  | App v 0.6.9 |  | No | Poornesh to check on this | 2026-08-14 | 2026-08-17
269 | closed | recording time is not showing total time when video is split into chunks | The recording time should show the total time that the recording has been elapsed. If I leave the device recording for >1hr, then it will split the video files into chunks, but when a new chunk starts, the time shown in the top border gets reset to 0:0:0. This is not right. It needs to keep showing the total accumulative time elapsed since recording was enabled, regardless of the file splits. | High | Application | Resolved | PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18 | Davis<6-July2026> Will trouble shoot and release in v 0.6.0 on 07/7 Davis<9-July2026> To be resolved in next release v0.6.1 on 7/14 Davis<14072026>implementation done in V0.6.2 to be released today <WW 16Jul2026>: Fixed in v0.6.2 | App: v0.6.2 |  | Yes
270 | closed | recording split files are 0bytes | when I enable recording for long periods of time (>1hr) and the video file gets split into chunks, the file size and contents of each chunk is 0Bytes | Low | Application | Resolved | PROVU_V_3.2.9_APP_v.0.5.9 PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | Davis<6-July2026> Will trouble shoot , post 07/7 release Davis<9-July2026> To be resolved in next release v0.6.1 on 7/14 Davis<14072026>implementation done in V0.6.2 to be released today Davis<16072026> Implementation & Testing done, v0.6.3 to be release either today else on 7/17 <WW 20Jul2026>: Still not resolved in v0.6.3. I see multiple files for a long recording period and several of the chunk files are 0Byte <Davis 22 JUl 2026> : Working on the bug , to be relsoved in v6.4 today's release. Added splitmuxsink, which is part of gstreamer. splitting is now handled in gstreamer. <WW 24Jul2026>: Tested on v0.6.4 and still not fixed. after recording continuous for 1hr:10min, I get the pop-ups showing recording will be split into multiple files, but when I stop the recording (after about 1hr 10min), I now only see one file in the media gallery and the file size is 0Bytes. <Davis 24Jul2026>: Found another issue when file size large during saving -fixed <WW 30Jul2026>: Tested in v0.6.6 and now resolved. | App v 0.6.5 |  | Yes
271 | closed | Shut down | sometimes when I have two cameras attached (broncho + USB) and I try to shut the device OFF (press and hold power btn for 2secs), it gets stuck in the 'shutting down' screen and does not shut off | High | Application | Resolved | PROVU_V_3.2.9_APP_v.0.5.9 | davis.packiaraj;#18/Poornesh | Poornesh/Devi<6-July2026> To release in Yacto v 3.3.0 on 07/7 <WW 16Jul2026>: Tested on v0.6.2 and have not seen this issue again. |  |  | Yes
272 | closed | viewport size and location wrong | For this issue refer to video : Demo_Video/v0.6.0 issues/viewport wrong and residual image after removal.mp4 1. I have a broncho and VL camera attached and showing in dual-view. 2. I select the bronchoscope image settings buttons and this brings up the left side panel with the image controls. 3. Then I remove the bronchoscope (do not press any buttons before this step). 4. The VL camera is shown in full view (ok) 5. Then re-attach the broncho within 5seconds 6. I see the broncho image squeezed into the RH side of the screen in the wrong place. VL camera view port does not change position. | High | Application | Resolved | PROVU_V_3.3.0_APP_v.0.6.0 | davis.packiaraj;#18 | Davis<9-July2026> To be resolved in next release v0.6.1 on 7/14 | App_v0.6.1 |  | Yes
273 | closed | audio tone incorrect when VL/VS camera is attached | The audio tone played when a USB camera is inserted is different to the bronchoscope. They should be the same? | High | Application | Resolved | PROVU_V_3.3.0_APP_v.0.6.0 | davis.packiaraj;#18 | Davis<9-July2026> To be resolved in next release v0.6.1 on 7/14 Davis<14072026>implementation done in V0.6.2 to be released today <WW 16Jul2026>: This was due to an erroneous fault being generated when a USB camera was inserted, which was causing a different audio tone to be played. This has been fixed in v0.6.2 | App: v0.6.2 |  | Yes
274 | CurrentPlan | reminder to re-enable error checking | some error have been masked for interim release for useability. Possibly software/hardware compatibility check and bronchoscope ID type. The issues with these checks must be resolved and re-enabled before launch | High | Application | Ready for testing | PROVU_V_3.3.0_APP_v.0.6.0 | davis.packiaraj;#18 |  | v0.7.7 |  | No | Davis to check on this | 2026-08-18
275 | closed | bronco LED switch-off not working from power-up | For special software for PISEO testing (v0.6.1), after power-up and log-in as Admin, I cannot select the broncho camera LED switch to OFF. I need to first select the LED-ON, then I can select the LED-OFF. When I do this, the LED can then be switched off, and subsequently switched back on etc. | High | Application | Resolved | PROVU_V_3.3.1_APP_v.0.6.1 |  | <WW 16 Jul2026> : Fixed in v0.6.2 | App: v0.6.2 |  | Yes
276 | closed | incorrect OV428 register 0x300e setting | Latest software for PISEO testing, OV428 register 0x300e is wrong. The register is reading back as 0xe7 but it needs to be 0xe5 | High | Application | Resolved | PROVU_V_3.3.1_APP_v.0.6.1 |  | <WW 16Jul2026>: Fixed in v0.6.2 | App: v0.6.2 |  | Yes
277 | closed | no audio tone on startup if HMDI monitor is attached | I do no hear an audio tone played on startup when the HDMI monitor is attached (from startup). If I remove all cameras from the display and remove the HDMI cable, just after startup, I hear the audio tone (device has powered up?). But if I power-up with HDMI monitor already attached, then the audio tone does not play | High | Application | Resolved | PROVU_V_3.3.2_APP_v.0.6.2 |  | <WW 24Jul2026>: Not seeing this in v0.6.4 |  |  | No
278 | closed | viewport size and location wrong - 2 | See video in shared TEAMS "../Demo Video/v0.6.3 issues/incorrect viewport and not showing dual view by default.mp4" If I have two cameras attached (it does not matter what the camera types are) and showing dual view. I cycle through removing one camera at a time, then both cameras, then re-attaching one camera at a time. In some situations (see video for example), the first viewport remains in the same place and size, but the second viewport is squeezed into the right-side. Note this is very similar to previous bug (now closed) #272. | High | Application | Resolved | PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | <Davis 22 JUl 2026> : Working on the bug , to be relsoved in v6.4 today's release <WW 24Jul2026>: Resolved in v0.6.4 | App v 0.6.4 |  | Yes
279 | closed | PLAY graphic does not completely cover the record icon during playback | the PLAY graphic does not completely cover the record icon during playback. Also the PLAY text is not perfectly centered on the white background | Low | Application | Resolved | PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 |  | App: 0.7.0 |  | No | Davis to check on this | 2026-08-14 | 2026-08-17
280 | closed | broncho image corruption when inserted very slowly | I can insert the bronchoscope plug very very slowly and this causes corruption issues and does not automatically recover or give error. To fix, I need to remove camera, then re-insert. I can insert the plug very,very slowly and there is a point where the software detects the presense of the camera, but does not show image frames. It does display the 'camera fault' screen, but then if I continue to push the plug all the way into the socket, I see corrupt images and it does not recover automatically | High | Application | No fix required | PROVU_V_3.3.3_APP_v.0.6.3 |  | <WW 22Jul2026>: When I repeat this, there are some (very infrequent) instances where the viewport is blank, but after a short period in this state the software correctly reports the 'Camera Error' notification. Closing this for now - tested on v0.6.3. |  |  | Yes
281 | closed | Admin account reset button feature not working | this small button under the battery tray cover does not seem to work. If any user account is in the settings menu and press+hold this button for >5sec, it should pop-up the admin account reset confirmation prompt | Low | Application | Resolved | PROVU_V_3.3.3_APP_v.0.6.3 | Poornesh/Devi/Davis | <WW 07Aug2026>: Ayush mentioned this only works if the touchpanel connection is missing!? <WW 22Sep2026>: Tested in v0.7.7 and resolved - BUT - it can be accessed from any screen, not just the settings screen. This is acceptable to us and is more intuitive. input docs and test methods will change to reflect this. | App v 0.7.2 |  | No | checking | 2026-08-14 | 2026-08-26
282 | closed | camera image adjustments are being changed when I change user accounts | The camera image adjustment controls are not being retained (if I change them) for each user account. Refer to video : "../Demo Video/v0.6.3 issues/camera image adjust not retained.mp4" | High | Application | Resolved | PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | <Davis 22 JUl 2026> : Working on the bug , to be relsoved in v6.4 today's release. This happens when color value is set to 10 - it resets to 0 - formatting issue fixed. <WW 24Jul2026>: for individual user accounts, this is resolved - can confirm that settings are now being correctly retained for each account - BUT - if I am logged in as an account and the settings have changed and then log-out into 'Guest' account, the previous settings are retained for the Guest account - is this correct? flexi to check requirements and revert back. <WW 30Jul2026>: Tested on v0.6.6 and resolved | App v 0.6.6 |  | Yes
283 | closed | can power down while it is going to sleep | During transition to sleep (after single press of power btn), I can press+hold the power button and the power-down option pop-up will appear. If I leave this, it will continue into sleep mode. Then when I wake it back up from sleep, this pop-up is still visible. This is a bit strange. Possiblt requirement clarification needed | Low | Application | Resolved | PROVU_V_3.3.3_APP_v.0.6.3 | davis, poornesh | <WW 17Aug2026>: retested in v0.7.0 and fixed. | App: 0.7.0 |  | No | Davis and Poornesh to check on this | 2026-08-14
284 | closed | can still use GUI controls during transition to sleep mode | during the transition to sleep mode, I can press buttons on the GUI and they still work - one example is the record button. I have been able to press the record button (video recording) just before it goes to sleep, then when it wakes up, the recording function is enabled (but not easy to reproduce). Suggest that all GUI controls are disabled when the user confirms the sleep mode | Low | Application | Resolved | PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | <WW 17Aug2026>: tested in v0.7.0 and fixed | App: 0.6.9 |  | No | Davis to check on this | 2026-08-14
285 | closed | bronchoscope image freeze | In some conditions of inserting and removing cameras, the bronchoscope image was frozen on the screen and no error was shown. Refer to demo video : '…/Demo Video/v0.6.3 issues/broncho freeze frame.mp4' | High | Application | Resolved | PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | <Davis 22 JUl 2026> : Working on the bug , to be relsoved in v6.4 today's release Related to #278 <WW 24Jul2026>: Resolved in v0.6.4 | App v 0.6.4 |  | Yes
286 | closed | EEPROM reading and writing to OV248 not working | the device needs to read the EEPROM from the bronchoscope every time a bronchoscope is inserted. If the EEPROM data is valid, it needs to send any OV428 register values to the OV428 - it needs to do this each time a bronchoscope is connected to the socket. | High | Application | Resolved | PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | <Davis 22 JUl 2026> : Working on the bug , to be relsoved in v6.4 today's release <WW 24Jul2026>: Retested on v0.6.4: I <think> I can program individual bronchoscope EEPROMs with data, and I see that the data is being retained on each broncho (but not checked independently). BUT the data registers for the OV428 are not being written to the OV428 chip when the broncho is attached. <WW 28Jul2026>: Tested and resolved in v0.6.5 | App: v0.6.5 |  | Yes
287 | closed | for VL/VS camera ports, cannot zoom in/out when in max viewport size | I connect two cameras to ports B and C. Port C is on the left viewport. Both are default to no zoom and normal viewport size. Now I select the max viewport size for PortB (bottom right hand size icon). This changes the icon to blue (all good) and shows the image in maximum viewport size (also good). But I cannot use the zoom in/out buttons. This needs to have zoom in/out feature on max viewport size. | High | Application | Resolved | PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | <Davis 22 JUl 2026> : Working on the bug , to be relsoved in v6.4 today's release. This change to be captured in SRS. <WW 24Jul2026>: Resolved in v0.6.4 | App v 0.6.4 |  | Yes
288 | closed | for VL/VS ports, zoom level is not obvious when switching between max viewport size | I have two cameras connected to portB and portC. PortC is on left viewport. Both viewports are normal size (max size button for each is grey). I now zoom in on PortB one level. It correctly shows the zoom image on PortB. Now I select the max viewport size for PortB and this fills the RHS viewport with the PortB image, but the image is zoomed in and I cannot use the zoom controls at the top of the window. The zoom level needs to be shown when in normal size viewport or max size viewport | High | Application | Resolved | PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | <Davis 22 JUl 2026> : As per the UI workflow, the zoom controls are disabled when in max view port.This change to be captured in SRS. <WW 24Jul2026>: Resolved in v0.6.4. SRS to be updated | App v 0.6.4 |  | Yes
289 | closed | For VL/VS ports, the viewport position changes if I close the RHS viewport, then re-enable streaming. | I have two cameras connected to PortB and portC. PortC is on the LHS. Both are default to no zoom and normal viewport size. Now I close the RHS viewport using the top 'X' control and it shows single viewport for portC - all good. Now I re-enable streaming of the PortB using the left hand panel control, and it shows the dual-view but shows the portC on the RHS - which is different to the previous side before I closed the port . This needs to be fixed so that the viewport positions are retained when streams are disabled/re-enabled without the camera being removed. | High | Application | Resolved | PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | <Davis 22 JUl 2026> : Flexicare to update the camera selection connection matrix to get this requirement clear. As per SRS -enabled camera will always take left view port. <Davis28JUL2026> Implementation in progress, to be released on 7/29 v6.7 Davis<08Aug2026>: change in requirement, flexicare to capture the new requirements for implementation <WW 17Aug2026>: retested in v0.7.0 and fixed. | App: 0.7.0 |  | Yes | ok for usability. For final - SRS and GTD 989 to be updated with new camera matrix . Davis to implement the same. | 2026-08-14
290 | closed | Frame-rate and latency is slower when 3 cameras are attached but only 2 streaming | I have 3 cameras physically attached - Bronchoscope, VL1 and VL2. Both VL are of the new type (ProVuLite). Logged in as Admin to monitor the frame-rate. PortB and PortC are streaming (dual-view). Broncho is physically attached, but not streaming. The frame-rate for each VL camera is reduced when the broncho is also physically attached. With the broncho physically attached (but not streaming), I am seeing VL fps down to 8-15fps which leads to noticable lag and stuttering. Without the bronchoscope attached (and two VL cameras attached and both streaming dual-view), fps increases to about 23fps which is expected for these VL cameras. (Even on single view for VL camera, the fps is about 25fps which is limited by the camera). | High | Application | Resolved | PROVU_V_3.3.3_APP_v.0.6.3 | davis.packiaraj;#18 | <Davis 23 JUl 2026> : To be released in v6.5 on 7/24 <Davis 27Jul2026>: Couldnt able to reproduce in my hardware, the fps when connecting 2 camera / 3 camera does not change. Devi to test with different hardware. <Davis 28Jul2026>: We were able to replicate the bug. This is happening when we attach bronco camera even with single USB camera. Main concern is HW ISP pipeline is handling the DMA in a non-polite way- Overloading DMA usage <Davis28JUL2026> As explaind discussed in today's call, PORT A to be disabled unless it is streamed /displayed <WW 06Aug2026>: Tested in v0.6.7 and not seeing this issue anymore. | App: v0.6.7 |  | Yes | ok
291 | closed | I am seeing more frequent instances of 'Camera ISP configuration failed ..' error | This seems to happen for various scenarious and does not seem to be repeatible. But simply removing and re-inserting the broncho causes this issue Refer to demo video: 'Demo Videos/v0.6.4 issues/Broncho failed ISP configuration.mp4' | High | Application | Resolved | PROVU_V_3.3.4_APP_v.0.6.4 | davis.packiaraj;#18 | Davis<27Jul2026>: i2c write to read is not synchronous. needed some time delay before reading after writing. Davis<30Jul2026>: Increasing timeout from 50 to 80ms Davis<06Aug26>: Removed timeouts between read/write. use i2c 5 bus for both read and write. <WW 06Aug2026>: Tested in v0.6.7 and not seeing this error. | App: v0.6.7 |  | Yes | testing ?, Davis to release v6.7 for testing
292 | closed | ShipMode from factory reset is not working | When we put the device into a factory reset, the device should put itself into 'ship mode' after it deletes all accounts and files. I see that it deletes the files, but after this it just resets and does not enter ship mode. Note that I tried this without a PSU attached and the device was running on battery. | High | Application | Resolved | PROVU_V_3.3.4_APP_v.0.6.4 | davis.packiaraj;#18 | <WW 22Sep2026>: Tested in version v0.7.7 and resolved. | App v0.7.2 |  | No | checking | 2026-08-25 | 2026-08-26
293 | closed | some language icons (flags) are missing from config pages | The country 'flag' icons are missing for Chinese, Arabic and Portugese. | High | Application | Resolved | PROVU_V_3.3.4_APP_v.0.6.4 | davis.packiaraj;#18 | Davis<27Jul2026>: added icons | App: v0.6.5 |  | Yes
294 | CurrentPlan | no recovery from auto-shut down in low battery state when charger is re-attached | IF the battery gets lower than 3%, then the battery enters shipping mode. BUT, if I plug the charger in when it shows the 'Device entering shipping mode..', then it does not clear this pop-up and does not show the charger attached. In addition; after this test, if I subseqently remove the charger again, the device does not enter ship mode | High | Application | Ready for testing | PROVU_V_3.3.4_APP_v.0.6.4 | davis.packiaraj;#18 |  | App v0.7.2 |  | No | Davis to check on this | 2026-08-25 | 2026-08-26
295 | CurrentPlan | battery lower level capacity threshold before entering shut-down | The lower level of battery capacity needs to change from 3% to 5%. This is based on recommendation from the battery supplier. DIDO and SRS to also be updated | High | Application | Ready for testing | PROVU_V_3.3.4_APP_v.0.6.4 | davis.packiaraj;#18 |  | App v 0.6.9 |  | No | Ayush-SRS to be updated, Davis to check on this. | 2026-08-11 | 2026-08-14
296 | CurrentPlan | OV428 default regs for vert line | going forward with new PCBs (V1.7 onwards), the default values for the following regs should be set if there is no EEPROM detected, or EEPROM values are corrupted: 1. 0x300e = 0xe6 2. 0x3601 = 0x80 | High | Application | Ready for testing | PROVU_V_3.3.5_APP_v.0.6.5 | davis.packiaraj;#18 | Davis<06Aug26>: Default 0x300e changed to 0xe6, 0x3601 is already 0x80 - no change required - Added in v0.6.7 This values will be applied irrespective of eeprom status. will this be ok? | App v 0.6.7 |  | No | Davis to check on this | 2026-08-11 | 2026-08-14
297 | CurrentPlan | Hardware PCB ID if invalid or not fitted | If the hardware PCB ID links/resistors are invalid, then the GUI presently does not show any number or value in the information page. This should be corrected to at least show 'Invalid' if an invalid combination is detected. SRS needs to be updated to provide guidance on this. | High | Application | Ready for testing | PROVU_V_3.3.5_APP_v.0.6.5 | davis.packiaraj;#18 |  | App v0.7.2 |  | No | Ayush-SRS to be updated, Davis to check on this. | 2026-08-18 | 2026-08-26
298 | closed | EEPROM Read/Write Latency | Time taken to read and send the data is much longer, 5 sec seen in v6.6. Actual time taken should be minimal as possible( <1sec) | High | Application | Resolved | PROVU_ V_3.3.6 & App v0.6.6 | davis.packiaraj;#18 | Davis<06Aug26>: Read and write now uses i2c 5 bus directly - significant improvement in time seen(v0.6.7) re-test <WW 06Aug2026>: Confirmed much faster time to adjust OV428 settings from when broncho is attached. Tested in v0.6.7 | App v 0.6.7 |  | Yes | release in v6.7, Davis to work on this
299 | closed | Issues coming out of sleep mode and inserting USB camera | The device can get into a constant cycle of reseting in the following scenario: 1. Attach all three cameras. Attach VS to portB and VL to portC. Enable streaming for portA and B only. 2. Remove portC camera. 3. put the device into sleep mode 4. Wake up the device and as soon as you see the blue background splash screen, quickly insert the portC camera. You should briefly hear some audio. 5. Now the FPGA emergency mode is shown, and sometimes the device will enter a state of constantly resetting. | High | Application | Resolved | PROVU_ V_3.3.7 & App v0.6.7 | davis.packiaraj;#18 | <WW 17Aug2026>: tested in v0.7.0 and fixed | PROVU_ V_3.3.8 & App v0.6.8 |  | Yes |  | 2026-08-07 | 2026-08-07
300 | CurrentPlan | Some hardware is not going into sleep mode | On one hardware device, we noticed that it was not entering sleep mode properly. The team found that the FPGA power enable line was not being de-asserted when entering sleep mode on this device. The result is that the screen is still on during sleep mode. | High | Multiple | In progress | PROVU_ V_3.3.7 & App v0.6.7 | Devi |  |  |  |  | Faulty device with v1.6 HW to be debug | NA | NA
301 | closed | Error Log not being Recorded | I check the USB copy error while copying the folder into the USB device, it raised the error for the overwriting the files. then i check the error log , this error was not listed. |  | Application | Resolved | PROVU_ V_3.3.8 & App v0.6.8 | davis.packiaraj;#18 | <WW 22Sep2026>: tested in v0.7.7. File over-writes are now handled and user prompt to select to overwrite. There is no error generated anymore. | App v0.7.2 |  |  |  |  | 2026-08-26
302 | closed | Gallery Folder Exporting failed | i trying duplicate folder to copy in the USB device , it raise an error. |  | Application | Resolved | PROVU_ V_3.3.8 & App v0.6.8 | davis.packiaraj;#18 | Duplicate of Bug#187 <WW 22Sep2026>: tested in v0.7.7 : pop-up now prompts user to overwrite existing files. I can successfully overwrite files. | App v0.7.2 |  |  |  |  | 2026-08-26
303 | closed | User Event Log exports the empty files | i have tryed exporting the events log , it gets successful message on device but when i check the file from PC it was empty. No log recorded |  | Application | Resolved | PROVU_ V_3.3.8 & App v0.6.8 | davis.packiaraj;#18 | <WW 17Aug2026>: tested on v0.7.0 and fixed. NOTE - this is the same as bug#173
304 | closed | every event seems to be appearing in a new pop-up at bottom of screen. | Most events seem to now be appearing as pop-ups on the GUI (bottom bar). E.g. when I attach a charger, I get a pop-up notification; when I remove the charger, I get a pop-up; when I delete all files in the gallery I get a pop-up. I did not think this was part of the requirements? I thought we only need to show pop-up notifications for those that are necessary, according to the table in the SRS? | High | Application | Resolved | PROVU_ V_3.4.0 & App v0.7.0 | davis.packiaraj;#18 | Davis 18Aug 2026-- Implemented as per latest SRS version B july 3rd dated. Davis <19Aug2026>: Updated as per new Temp SRS Event table <WW 20Aug2026>: retested in v0.7.1 and is now fixed. | App v 0.7.1 |  | Yes |  | 2026-08-18 | 2026-08-19
305 | closed | USB stick does not get recognised when I wake from sleep | The device does not recognise a USB attached/inserted when it wakes from sleep. | High | Multiple | Resolved | PROVU_ V_3.4.0 & App v0.7.0 | Poornesh/Devi | Davis 18Aug 2026--To be resolved in v7.1 Poornesh/Devi 19Aug 2026- Not able to reproduce the bug <WW 20Aug2026>: I tested again on v0.7.1 and am not seeing this issue. |  |  | Yes |  | 2026-08-18
306 | CurrentPlan | I am seeing a high number of USB host#1 and host#2 supply rail invalid errors | I see a lot of new errors in the log for 'USB Host#1 supply rail invalid' and 'USB Host#2 supply rail invalid'. Curiously, I do not see the error pop-ups on the GUI. Is this because some errors have been masked for ux study software verison? There is nothing wrong with the hardware because the cameras work well. | High | Multiple | Ready for testing | PROVU_ V_3.4.0 & App v0.7.0 | davis.packiaraj;#18 | Davis<19Aug2026>: Masked this error for now in v0.7.1 Davis<23Sep2026>: Unmasked in v7.7 | App v0.7.7 |  | No |  | 2026-08-18 | 2026-08-19
307 | CurrentPlan | PortB USB camera insertion, I see device get reset | Connect the USB camera to port#B and wait till device switch to dual view port. remove the camera before device start the feed in view port. the device will get reset | High | Multiple | Ready for testing | PROVU_ V_3.4.0 & App v0.7.0 |  | Poornesh<19Aug2026>: As informed in todays syncup call , when we try the same way Ayush is doing (frequently removing and insering without wating for feed) , I observed that the crash is occurring within the UVC driver path itself. This requires deeper debugging and analysis There will be no issue if we wait for a second then remove the camera and reinsert for multiple times . | App v 0.7.2 |  |  |  |  | 2026-08-26
308 | CurrentPlan | Power cycle the device and you will see "internal system error" | Shutdown the device from the GUI, pressing power button and selecting shutdown prompt. Power up it again and you will "error pop-up, internal system error ... " it appears as a flash, come and go in a second I do not have the battery connected in my device | High | Multiple | Ready for testing | PROVU_ V_3.4.0 & App v0.7.0 | davis.packiaraj;#18 | Davis 19Aug 2026- As per 19/8 Sync up call, Since UX testing is done with battery , as of now no issue. However need to resolve Raise error only after 2 sec for process errors | App v 0.7.2 |  | Yes |  |  | 2026-08-26
309 | CurrentPlan | USB camera Reset multiple times (power up) | I powered up device with bronchoscope plugged in, then connected the USB camera (VS) on portB using AMP cable. - Initially, camera LED taking blinking and device continue in a loop of camera connected/disconnected - After few iteration (aprox 30 seconds, the LED blink stops and no camera feed on display. Also, device showing empty viewport without any error. | High | Multiple | Ready for testing | PROVU_ V_3.4.0 & App v0.7.0 | Poornesh/Devi |  | App v 0.7.2 |  | Yes |  |  | 2026-08-26
310 | CurrentPlan | USB camera Reset multiple times (sleep) | I have connected the USB camera wih device, then pressing the power button I put the device in sleep mode. Disconnected the camera while device is in sleep, then wakeup device from sleep. Then plug the USB camera again and it repeatedly showing connect/disconnect. | High | Multiple | Ready for testing | PROVU_ V_3.4.0 & App v0.7.0 | Poornesh/Devi |  | App v 0.7.2 |  | Yes |  |  | 2026-08-26
311 | CurrentPlan | Device goes from sleep to power cycle | If I do not connect the battery to device, and put device in sleep mode. Later after 2 min, when trying to wakeup device from sleep, it restarts. I have confirmed that this condition does not occur with battery | Low | Yocto | Not able to reproduce | PROVU_ V_3.4.0 & App v0.7.0 | Poornesh |  |  |  | No
312 | changeRequirement | Storage space remaining displayed to align with total eMMC size | In the information page (and any other related pages), the remaining storage space displayed should change to the format : "xxGB used of 32GB". The 'xx' value will represent the total used amount including any apps, backup data, settings, media files, etc (similar to modern smart phones). The 32GB will be fixed for our device because the eMMC is fixed at 32GB. This should align with the IFU because we state our total storage space in the IFU as 32GB. | Low | Multiple | Require inputs from Flexi | PROVU_ V_3.4.0 & App v0.7.0 | davis.packiaraj;#18 | <Davis 15 sep2026>: To Discuss in today's Sync up call. New GUI screen update+ Translation update
313 | changeRequirement | Broncho Gamma value to change | The default gamma value for bronchoscope (for all accouts) needs to be changed to 50. CONFIRMED | High | Application | Ready for testing | PROVU_ V_3.4.1 & App v0.7.1 | davis.packiaraj;#18 |  | App v0.7.2 |  |  |  |  | 2026-08-26
314 | CurrentPlan | Admin account was not created properly during setup wizard | recent Ux team found this issue: On factory reset, the device reboots on completion (i.e. not triggering ship mode), and then goes into set-up wizard. If the wizard is not completed and the device turns off, turning it back on does not always retrigger the set-up wizard. Instead it boots normally, which then prevents you from doing anything account wise or from being able to re-reset the device. Ayush advised that the fix for this was to use fixed temporary credentials (Admin/fcadmin123) to access admin account. This is not acceptable for launch and needs to be considered. User must always be able to access or create a Admin account, or home screen must not be available unless a admin account has been created. | High | Application | Ready for testing | PROVU_ V_3.4.2 & App v0.7.2 | davis.packiaraj;#18 | <WW 10Sep2026>: Video demonstration of this found here : TEAMS shared drive/Demo_Video/v0.7.2 issues/Set-Up Wizard Sleep Bug.mp4' <Davis 15 sep2026>: To be corrected in today's release v7.7 dated 15/9 | App v0.7.7
315 | changeRequirement | Text string for 'Colour' needs to change for American translation | Feedback from ux study (Aug2026) : For the Americans, the word 'colour' needs to be changed to 'color' in all occurances | High | Application | Ready for testing | PROVU_ V_3.4.1 & App v0.7.1 | davis.packiaraj;#18 |  | App: v0.8.0
316 | changeRequirement | Text string for 'forename' and 'surname' to change to 'first name' and 'last name' for American only | Feedback from ux study (Aug2026) : For the Americans, the words 'forename' and 'surname' needs to be changed to 'first name' and 'last name' repectively in all occurances | High | Application | Ready for testing | PROVU_ V_3.4.1 & App v0.7.1 | davis.packiaraj;#18 |  | App: v0.8.0
317 | CurrentPlan | Long delay when selecting Factory Reset from the GUI | When we put the device into a factory reset, there is a very long delay when the 'Yes' button is pressed. One tester also found that if the device is forced to OFF state (using power button) during this wait period, when it is re-powered, there was a battery error alert | High | Application | Ready for testing | PROVU_ V_3.4.2 & App v0.7.2 | davis.packiaraj;#18 | <Davis 15 sep2026>: To Discuss in today's Sync up call. New GUI screen from Abhi, Factory reset in progress... This popup will be added now during factory reset | App: v0.8.0
318 | Backlog | FPGA status checking to be provided to application layer by driver layer | The linux driver layer needs to determine when the FPGA is ready to be used by the application and set some flag or status that can be read by the application. This will simplify the architecture and provide a more robust to communicate with the FPGA and reduce the likelihood of corruption. | High | Multiple | Unassigned | PROVU_ V_3.4.7 & App v0.7.7
319 | Backlog | FPS data shown on the viewport to be removed for Admin users | The fps data presently shown on the viewport needs to be removed for all users. Needs to be removed for admin users. | High | Application | Ready for testing | PROVU_ V_3.4.7 & App v0.7.7 |  |  | App: v0.8.0
320 | Backlog | LED-OFF selection control in the admin user image adjustment menu to be removed | The LED-OFF selection needs to be removed for the admin user account image adjustment menu | High | Application | Ready for testing | PROVU_ V_3.4.7 & App v0.7.7 |  |  | App: v0.8.0
321 | Backlog | Any non-essential image control adjustments to be removed from Admin user account | controls like black level adjust, priority ordering of the controls, sharpness, etc, to be removed. SRS needs to cleraly define what is contained | High | Application | Ready for testing | PROVU_ V_3.4.7 & App v0.7.7 |  |  | App: v0.8.0
321 | Backlog | navigation icon issue in gallery screen | If I go into the media gallery and keep deleting files, it gets to the last file and the left hand navigation arrow is still showing although it does not do anything | High | Application | Ready for testing | PROVU_ V_3.4.7 & App v0.7.7 |  |  | App: v0.8.2
322 | CurrentPlan | I am seeing frequent 'Battery fault detected' faults | I am seeing frequent 'Battery fault detected' faults appearing. The batteries are good. | High | Application | Ready for testing | PROVU_ V_3.4.7 & App v0.7.7 |  | <WW 23Sep2026>: To be fixed for next release. Register 0x00 will be corrected to 0x26 (set and expected value). But Reg0x06 will be masked for this version. Flexicare needs to investigate this further before confirming the value. <Davis 24Sep2026>: Maked and v7.9 released | v7.9
323 | Backlog | I am not seeing the message bar scroll through the faults if more than 1 fault has been raised. Is this a requirement? | when multiple faults have occurred, should the messages be scrolling through (after a fixed time period to show each message) on the bottom bar? If not, how will the user see all the faults from the home screen? Only admin can access the logs? | High | Application | Ready for testing | PROVU_ V_3.4.7 & App v0.7.7 |  |  | App: v0.8.0
324 | CurrentPlan | All data entry fields to have a maximum fixed length limit and checks | All user inputs shall have a maximum input length check <WW 24Sep2026>: Limit ALL text entry fields to 15chars max | High | Application | Ready for testing | PROVU_ V_3.4.9 & App v0.7.9 |  | <Davis 24Sep26>: Currently character limit checks applied to user name, fname, sname, password to 30 char as per the SRS. Does that stays the same or do we need to limit to 15 ? This check is missed in update user account, which is added in v0.8.0. As discussed with Wayne, characters will be limitted to 30 | App: v0.8.1
325 | CurrentPlan | I can not access the Error or Event log on my device | In the software version application 7.9, when I click over the button Error or Event then it opens a dialog saying "Do you want to cancel? ". It does not open error or event logs | High | Application | Ready for testing | PROVU_ V_3.4.9 & App v0.7.9 |  |  | App: v0.8.0
326 | CurrentPlan | Device does not go to the Power OFF when recording is ON | during the power-off operation when recording is ON, the device displays the bronchoscope emergency-mode camera feed; the screen fluctuates multiple times, and the device restarts instead of entering power-off mode. (Device not Powering off with power Button press if recording is enabled) | High | Application | Ready for testing | PROVU_V_3_5_1_APP_0_8_1 | davis.packiaraj;#18 |  | App: v0.8.2
327 | CurrentPlan | Device restart over software upgrade page | Device restarts over software upgrade page when battery is not connected. most of the time it happens when software is scanning for USB devices | High | Application | Not able to reproduce | PROVU_V_3_5_1_APP_0_8_1 | davis.packiaraj;#18 | <Davis 29Sep2026>: Tried replicating the behaviour, by disconnecting the battery and running upgrade for atleast 10 times, couldn't reproduce. Is there a sequence we need to follow?
328 | CurrentPlan | External device Like mouse can work with USBA, MicroUSB ports | I plugged in mouse on the USBA port and MicroUSB port. I can access the device and control through mouse | High | Yocto | Ready for testing | PROVU_V_3_5_1_APP_0_8_1 |  |  | App: v0.8.2
329 | Backlog | Device boot failed after upgrade | Tried 6 devices to load with version 3.4.9, app7.9 production version device failed to boot if charger is plugged in. with battery device is fine. The display stuck at progress bar | High | Multiple | In progress | PROVU_ V_3.4.9 & App v0.7.9 (prodcution) | Poornesh/Davis | <Poornesh29sep2026>: Trying to replicate in Mindteck devices
330 | changeRequirement | Database migration issue - Device Reset | Device resets after software upgrade on Account creation. Database migration issue CRC added to the database which changed the database schema. New software expects new schema (current fix: delete the old database and re-launch) Proposed: Check the schema and fail login, so that crash is avoided | High | Application | Ready for testing | PROVU_ V_3.4.9 & App v0.7.9 (prodcution) | davis.packiaraj;#18 | <Davis 29Sep2026>: If database schema does not match, login will fail. For recovery from this situation, shall we use admin reset button to delete the user accounts database instead of just deleting the admin user accounts?
331 | changeRequirement | Enable HW watch dog during the boot process | The software shall enable the hardware watchdog timer during the boot process. If the boot process stuck, the watchdog shall reset the device. Upon reboot, the boot script shall increment the boot counter. If the boot counter exceeds the configured threshold, the device shall initiate recovery using the secondary partition | High | Yocto | In progress | PROVU_ V_3.4.9 & App v0.7.9 (prodcution) | Poornesh/Devi |  | App: v0.8.2
332 | CurrentPlan | Factory reset issue on camera view port | from a factory reset, when I attach two cameras (Bronc + VL in portC), it does not default to dual view | High | Application | Not able to reproduce | PROVU_ V_3.4.9 & App v0.7.9 | davis.packiaraj;#18
333 | Backlog | Battery max current value register to be checked | This is related to bug#322. The software should be able to check that the value for max current (on input side) is correct from the BQ charger chip; on an earlier release, this value was not as expected. Flexi needs to find out how this value is updated (within the BQ charger chip) and make sure it is read correctly. This register is presently being masked by software, but this needs to be re-enabled when the mechanism is understood | High | Application | Unassigned | PROVU_ V_3.4.9 & App v0.7.9
334 | Backlog | Watchdog reset feature improvements | The watcdog reset improvments like enabling from uboot level is masked for now. The following requirement considered for round 2 updates - Attempt three retries to boot a partition and if all three partition failed then boot from secondary partition | High | Yocto | Unassigned
335 | Backlog | Application integrity check | The application integrity check and booting from alternative partition over failure is masked and need graphics updated. It is considered to be updated in round 2 | High | Yocto | Unassigned
336 | Backlog | Image Controls | The Adming and standard user image controls will be same as Brightness, Contrast, Satuartion, Color (R+B) | High | Documentation | In progress | n/a |  | Ayush: Need to update SRS
337 | CurrentPlan | View Port blinks for second camera connections | When the second camera is connected the feed area of the screen got blank for a second then both feed gets displayed. | Low | Application | Require inputs from Flexi | PROVU_V_3_5_2_APP_0_8_2 | Abhishek | Abhi: As per davis, this can not be fix as we are using the single view port to display the camera feeds.
338 | CurrentPlan | Error Shown:- Fault Detected On camera Port A ID-00034 | This error is comming when bronchoscope camera is disabled, as soon camera feed gets active it goes off. Broncoscope analog ID invalid | High | Application | Ready for testing | PROVU_V_3_5_3_APP_0_8_3 | davis.packiaraj;#18 |  | App v0.8.4
339 | CurrentPlan | Device getting reset when multiple tap on settings and gallery icon from right side panel | From the settings page when tap gallery button then agin settings button , Repeating this to multiple time quickly device gets reset. | High | Application | Ready for testing | PROVU_V_3_5_3_APP_0_8_3 | davis.packiaraj;#18 |  | App v0.8.4
340 | CurrentPlan | Device Reset (transition to sleep) | 1. Case1: if we remove USB camera while device going to sleep, the device restarts 2. Case2: if we plug all three camera(s), power adapter and put device into sleep mode, sometime device restart instead of going to sleep mode | High | Application | Ready for testing | PROVU_V_3_5_3_APP_0_8_3 | davis.packiaraj;#18 |  | App v0.8.4
341 | CurrentPlan | Sleep Mode (device stuck in emergecy mode) |  | Low | Application | Resolved | PROVU_V_3_5_2_APP_0_8_2 | davis.packiaraj;#18 |  | App v8.3
342 | CurrentPlan | Device not Entering Shipping Mode while battery < 5% | 1. Connect USB and broncoscope camera and place device into sleep mode. 2. Disconnect USB camera and wakeup device from sleep. Sometime device stuck in emergecny mode | High | Application | Ready for testing | PROVU_V_3_5_3_APP_0_8_3 | davis.packiaraj;#18 |  | App v0.8.4
343 | CurrentPlan | Device auto power OFF from sleep mode | Device does not go to auto power OFF from Sleep mode (20 min) | Low | Application | Require inputs from Flexi | PROVU_V_3_5_3_APP_0_8_3 | WW | Abhi: as per davis, This is not feasible from application level.