## CoverPage
155 - ProVu Accessories
ProVu Vista Display – Software Integration Test Request - 01
Project: |  | 155 - ProVu Accessories
Document ID/Version: |  | GTD-001224/A
DOCUMENT HISTORY
Version | Document Changes |  | Author | Date
A | Initial release |  | Abhishek Shukla | 2026-09-16
Associated Documents
Document No |  | Document Title
SRS-000053 |  | ProVu Vista Display –SOFTWARE REQUIREMENT SPECIFICATION
GTD-000993 |  | PROVU2 LARGE DISPLAY Software Test Plan
Attachment
Name |  | Description And Summary
Attachments |  | Attachement having test binaries
Note: Use the Updated test binaries from Azure on given link : https://dev.azure.com/FlexicareDevOps/ProVu%2015/_git/provu2_test_suite

## About
GTD-001224: Software Integration Test Request-01
Test Request Information
Author | Abhishek Shukla
Revision | xxx | Software Version | STP
Reference | GTD-000993 | Hardware Version | STP | Date | 2026-09-16
Test objective
Verify the ProVu Vista Display camera integration functions covering camera image acquisition, image processing, display output, camera detection, camera power control, camera LED brightness, and camera image property handling as per the associated software requirements.
Test Coverage
Test coverage areas: • Camera detection, connection/disconnection, and camera-type compatibility • Supported video formats, resolution, aspect ratio, frame rate, and latency • Image exposure, zoom, orientation, and display output • LCD brightness, Dual View operation, and screen recording • Camera frame-loss detection, error handling, and audible notifications • Camera power control, Power-Good GPIOs, and RK3568-to-FPGA I2C configuration • USB camera detection and power control • Bronchoscope EEPROM detection, read/write access, and OV428 register configuration SRS requirements covered (26): Camera: SRS-FUN-CAM-010, SRS-FUN-CAM-020, SRS-FUN-CAM-030, SRS-FUN-CAM-040, SRS-FUN-CAM-050, SRS-FUN-CAM-060, SRS-FUN-CAM-070, SRS-FUN-CAM-080, SRS-FUN-CAM-1460, SRS-FUN-CAM-1680-01, SRS-FUN-CAM-1680-07, SRS-FUN-CAM-1690, SRS-FUN-CAM-1700, SRS-FUN-CAM-1710, SRS-FUN-CAM-1820, SRS-FUN-CAM-1840, SRS-FUN-CAM-2030, SRS-FUN-CAM-2080, SRS-FUN-CAM-2170 GUI: SRS-FUN-GUI-210, SRS-FUN-GUI-220, SRS-FUN-GUI-270 USB: SRS-FUN-USB-1660, SRS-FUN-USB-1670 Power: SRS-FUN-PWR-1890 Software: SRS-FUN-SFT-200
Test Execution Instructions
1. Execute the camera integration test cases in the sequence defined in the "Test Script" sheet. Ensure the required cameras, LED blinker/LDR setup, logic analyzer, and device display are connected where applicable. 2. Record all measured latency, ADC, detection time, brightness, and image-property results in the "Actual Result" column.
Test environment
Production-complete ProVu Vista Display device powered by its battery, with the device serial port connected to a Windows/Linux PC for running commands and verifying command output and logs. A logic analyzer used to test and verify camera and control signals. Supported cameras: • VS camera • VL camera • Bronchoscope camera
Definitions or Acronyms Used
VS – Video Stylet VL – Video Laryngoscope ADC – Analog-to-Digital Converter GPIO – General Purpose Input/Output FPGA – Field Programmable Gate Array
Instruments & Tools
• Production-complete ProVu Vista Display device with battery • Supported VS, VL, and bronchoscope cameras • Bronchoscope development board with camera • USB-to-serial/serial-port cable and required connection cables • Windows/Linux PC with a terminal application • Logic analyzer and PicoScope • Device terminal/log interface and screen-recording function
Fail/Pass Criteria
The test will be considered as pass if actual result meets the expected criteria in sheet "Test Script"
Test Report Reference:
/* Add the related Test Report document number here */
Comments:
Note: 1. After each test completion of given test IDs Power OFF the device and remove the camera cables , and on startup of next test follow the instructions. 2. For serial cable connection sniff the Tx, Rx and GND for serial communication from the device, Make sure the software used has the enabled serial cable. 3. Used $root for login in serial and dmesg -n 2 for stopping the continue message. Note: Use the Updated test binaries from Azure on given link : https://dev.azure.com/FlexicareDevOps/ProVu%2015/_git/provu2_test_suite before starting the tests of this scripts. Note: The software Being tested here is serial enabled version.

## Test Script
GTD-001224: Software Integration Test Request-01
Test Case No. | Test Descriptions | Expected Result | Actual Result
The below steps verify the software version
10 | Power On the device and configured the welcome page. | Setup | Setup
20 | Go to device information page | Setup | Setup
30 | Verify that the device displays the following version information on the screen: 1. OS Build Version 2. Application Version 3. Hardware Version 4. FPGA Version | Device should display the version information correctly. | Does the device display the version information correctly? _____ Record the versions: OS Build Version: _____ Application Version: _____ Hardware Version: _____ FPGA Version
40 | Verify that the displayed versions match the versions mentioned on the "About" page. | Device version information should match the "About" page. | Does Device version information should match the "About" page?_____
Test ID: STS-ProvUVista-Integration-000010 Software item (under test): SRS-FUN-CAM-010 , SRS-FUN-CAM-020 , SRS-FUN-CAM-030
50 | Connect the VL camera to the ProVu Vista Display device on PORT B and power ON the device. Connect the device serial port to a Windows/Linux PC and open the terminal interface. Run command: v4l2-ctl --list-devices | Device should boot successfully and the VL camera should be detected by the terminal command output. Expected output should show: VL Camera /dev/video9 | Is the VL camera connected and detected? _____ Detected device node: _____ Does device display the camera feed?___
60 | Check the supported VL camera formats and current active video format using the terminal commands. Check supported formats: v4l2-ctl -d /dev/video9 --list-formats-ext Check current format: v4l2-ctl -d /dev/video9 --get-fmt-video Verify The Aspect ratio for camera Port B from sheet "Aspect ratio Test" Table 1 | Terminal shall display the Supported format output should include: Pixel Format: 'YUYV' Size: Discrete 640x480 Current format output should show: Width: 640 Height: 480 Pixel Format: YUYV Device pass the spect ration test for camera Port B from sheet "Aspect ratio Test" Table 1. | Does Supported format contains YUYV 640x480? _____ Recorded current format: Width: _____ Height: _____ Pixel Format: _____ Is current format YUYV and resolution 640 × 480? _____ Does device Pass the aspect ratio Test from for camera Port B from sheet "Aspect ratio Test" Table 1 ?___
70 | Remove the Vl camera from Port B. | Setup | 155 - ProVu Accessories
80 | Connect the VL camera to the ProVu Vista Display device on PORT C Run command: v4l2-ctl --list-devices | The VL camera should be detected by the terminal command output. Expected output should show: VL Camera /dev/video9 | Is the VL camera connected and detected? _____ Detected device node: _____ Does device display the camera feed?___
90 | Check the supported VL camera formats and current active video format using the terminal commands. Check supported formats: v4l2-ctl -d /dev/video9 --list-formats-ext Check current format: v4l2-ctl -d /dev/video9 --get-fmt-video Verify The Aspect ratio for camera Port C from sheet "Aspect ratio Test" Table 2 | Terminal shall display the Supported format output should include: Pixel Format: 'YUYV' Size: Discrete 640x480 Current format output should show: Width: 640 Height: 480 Pixel Format: YUYV Device pass the spect ratio test for camera Port C from sheet "Aspect ratio Test" Table 2 | Does Supported format contains YUYV 640x480? _____ Recorded current format: Width: _____ Height: _____ Pixel Format: _____ Is current format YUYV and resolution 640 × 480? _____ Does device Pass the aspect ratio Test from for camera Port C from sheet "Aspect ratio Test" Table 2 ?___
100 | Remove the Vl camera from Port C. | Setup | Setup
110 | Connect a Webcam on Device Port B which has the resolution of 1920*1080. Check the supported VL camera formats and current active video format using the terminal commands. Check current format: v4l2-ctl -d /dev/video9 --get-fmt-video Check the device screen for camera error and warning message. Navigate to the device error log page. Verify the ERROR log ID00030. | Device Terminal log shall display the supported resolution as 1920*1080. Device should display the Camrea Error screen on display. Device shall display the Error text as "Fault detected on camera Port 2. Check camera and connections." Device shall display the Error log ID: 00030 with below texts "Fault with camera port#2 ( Accessories )" | Recorded current format: Width: _____ Height: _____ Pixel Format: _____ Is mismatch resolution 1920 × 1080 observed? _____ Is Camera Error screen displayed? _____ Displayed error text: "Fault detected on camera Port 2. Check camera and connections." _____ Error log ID recorded: _____ Is Error ID 00030 logged? _____ Error log text: "Fault with camera port#2 (Accessories)" _____
120 | Remove the Webcam camera from Port B. | Setup | Setup
130 | Connect a Webcam on Device Port C which has the resolution of 1920*1080. Check the supported VL camera formats and current active video format using the terminal commands. Check current format: v4l2-ctl -d /dev/video9 --get-fmt-video Check the device screen for camera error and warning message. Navigate to the device error log page. Verify the ERROR log ID00031. | Device Terminal log shall display the supported resolution as 1920*1080. Device should display the Camrea Error screen on display. Device shall display the Error text as "Fault detected on camera Port 3. Check camera and connections." Device shall display the Error log ID: 00031 with below texts "Fault with camera port#3 ( Accessories )" | Recorded current format: Width: _____ Height: _____ Pixel Format: _____ Is mismatch resolution 1920 × 1080 observed? _____ Is Camera Error screen displayed? _____ Displayed error text: "Fault detected on camera Port 3. Check camera and connections." _____ Error log ID recorded: _____ Is Error ID 00031 logged? _____ Error log text: "Fault with camera port#3 (Accessories)" _____
140 | Remove the Webcam camera from Port C. | Setup | Setup
150 | Connect the VS camera to the ProVu Vista Display device on PORT B Run command: v4l2-ctl --list-devices | VS camera should be detected by the terminal command output. Expected output should show: VL Camera /dev/video9 | Is the VS camera connected and detected? _____ Detected device node: _____ Does device display the camera feed?___
160 | Check the supported VS camera formats and current active video format using the terminal commands. Check supported formats: v4l2-ctl -d /dev/video9 --list-formats-ext Check current format: v4l2-ctl -d /dev/video9 --get-fmt-video | Terminal shall display the Supported format output should include: Pixel Format: 'YUYV' Size: Discrete 640x480 Current format output should show: Width: 640 Height: 480 Pixel Format: YUYV | Does Supported format contains YUYV 640x480? _____ Recorded current format: Width: _____ Height: _____ Pixel Format: _____ Is current format YUYV and resolution 640 × 480? _____
170 | Remove the VS camera from Port B. | Setup | Setup
180 | Connect the VS camera to the ProVu Vista Display device on PORT C Run command: v4l2-ctl --list-devices | The VS camera should be detected by the terminal command output. Expected output should show: VL Camera /dev/video9 | Is the VS camera connected and detected? _____ Detected device node: _____ Does device display the camera feed ?___
190 | Check the supported VL camera formats and current active video format using the terminal commands. Check supported formats: v4l2-ctl -d /dev/video9 --list-formats-ext Check current format: v4l2-ctl -d /dev/video9 --get-fmt-video | Terminal shall display the Supported format output should include: Pixel Format: 'YUYV' Size: Discrete 640x480 Current format output should show: Width: 640 Height: 480 Pixel Format: YUYV | Does Supported format contains YUYV 640x480? _____ Recorded current format: Width: _____ Height: _____ Pixel Format: _____ Is current format YUYV and resolution 640 × 480? _____
200 | Remove the VS camera from Port C. | Setup | Setup
210 | Connect the Bronchoscope camera to the ProVu Vista Display device on PORT A Run command: v4l2-ctl --list-devices | The VS camera should be detected by the terminal command output. Expected output should show: /dev/video0 | Is the Bronchoscope camera connected and detected? _____ Detected device node: _____ Does device display the camera feed ?___
220 | Check the supported Bronchoscope camera formats and current active video format using the terminal commands. Check supported formats: v4l2-ctl -d /dev/video0 --list-formats-ext Check current format: v4l2-ctl -d /dev/video0 --get-fmt-video Verify The Aspect ratio for camera Port A from sheet "Aspect ratio Test" Table 3 | Terminal shall display the Supported format output should include: Pixel Format: "UYVY' Size: Discrete 400x400 Current format output should show: Width: 400 Height: 400 Pixel Format: UYVY Device pass the spect ration test for camera Port A from sheet "Aspect ratio Test" Table 3. | Does Supported format contains UYVY 400x400? _____ Recorded current format: Width: _____ Height: _____ Pixel Format: _____ Is current format UYVY and resolution 400 × 400? _____ Does device Pass the aspect ratio Test from for camera Port A from sheet "Aspect ratio Test" Table 3 ?___
230 | Remove the Bronchoscope camera from Port A. | Setup | Setup
240 | Connect the faulty Bronchoscope. Verify the camera Error screen and warning message on the proVu device Navigate to the error log screen and verify the error log message and ID. Note: For creating the faulty bronchoscope use the Bronchoscope dev board and remove the camera while keep the board connected. | Device should display the Camera Error screen on display. Device shall display the Error text as "Fault detected on camera Port 1. Check camera and connections." Device shall display the Error log ID: 00029 with below texts "Fault with camera port#1 ( Accessories )" | Is Camera Error screen displayed? _____ Displayed error text: "Fault detected on camera Port 1. Check camera and connections." _____ Error log ID recorded: _____ Is Error ID 00029 logged? _____ Error log text: "Fault with camera port#1 (Accessories)" _____
250 | Remove the Bronchoscope camera from Port A. | Setup | Setup
Test ID: STS-ProvUVista-Integration-000020 Software item (under test): SRS-FUN-CAM-040
260 | Connect the VS camera to Port B, the VL camera to Port C, and the Bronchoscope camera to Port A. Activate the VS and Bronchoscope camera feeds in Dual View mode. | The device shall detect all three cameras and display the VS and Bronchoscope camera feeds simultaneously in Dual View mode. | Are all three cameras detected? _____ Are the VS and Bronchoscope feeds displayed in Dual View mode? _____
270 | Start screen recording on the device while the VS and Bronchoscope camera feeds are active. | Screen recording shall start successfully without interrupting either active camera feed. | Does screen recording start successfully? _____ Do both active camera feeds remain uninterrupted? _____
280 | Place a stopwatch that displays milliseconds beside the ProVu Vista Display. Position the VS camera so that it clearly captures the stopwatch, and verify that the stopwatch time is visible on the device screen. | The VS camera feed shall clearly display the stopwatch time, including milliseconds, on the device screen. | Is the stopwatch clearly visible in the VS camera feed? _____ Are the milliseconds readable on the device screen? _____
290 | Start the stopwatch. Using an external camera or mobile phone, capture the physical stopwatch and the ProVu Vista Display in the same image. Capture 10 images at 5-second intervals. | Ten images shall be captured at approximately 5-second intervals. Each image shall clearly show both the physical stopwatch time and the corresponding time displayed on the device screen. | Number of images captured: _____ Were the images captured at approximately 5-second intervals? _____ Does each image clearly show both times? _____
300 | For each of the 10 captured images, enter the physical stopwatch time and the corresponding time shown on the device display in the "Latency Test Record" sheet. Verify that the sheet calculates and records the latency for each image. | The physical and displayed stopwatch times shall be entered for all 10 images. The "Latency Test Record" sheet shall calculate and display the latency for each image in milliseconds. | Are both stopwatch times entered for all 10 images? _____ Does the "Latency Test Record" sheet calculate all 10 latency values? _____
310 | Record the average latency calculated in cell E23 of the "Latency Test Record" sheet. Do not calculate the average manually. | Cell E23 of the "Latency Test Record" sheet shall calculate the average from the recorded latency values. The calculated average shall be recorded in the Actual Result. | Average latency recorded from the "Latency Test Record" sheet, cell E23: _____ ms
320 | Verify that the average latency calculated in cell E23 of the "Latency Test Record" sheet is less than 180 ms. | The average latency calculated in cell E23 of the "Latency Test Record" sheet shall be less than 180 ms. | Calculated average latency from cell E23: _____ ms Is the calculated average latency less than 180 ms? _____
330 | Deactivate the Bronchoscope camera feed so that only the VS camera feed is active. Repeat Steps 280–320 and record the VS camera latency in Single View mode. | The latency measurement shall be completed with only the VS camera feed active, and the calculated average latency shall be less than 180 ms. | VS camera average latency in Single View mode: _____ ms Is the average latency less than 180 ms? _____
340 | Activate the VS camera feed on Port B and the VL camera feed on Port C in Dual View mode. Repeat Steps 280–320 with the VL camera focused on the stopwatch, and record the VL camera latency. | The latency measurement shall be completed for the VL camera while the VS and VL feeds are active in Dual View mode. The calculated average latency shall be less than 180 ms. | VL camera average latency in Dual View mode: _____ ms Is the average latency less than 180 ms? _____
350 | Activate the VS camera feed on Port B and the Bronchoscope camera feed on Port A in Dual View mode. Repeat Steps 280–320 with the Bronchoscope camera focused on the stopwatch, and record the Bronchoscope camera latency. | The latency measurement shall be completed for the Bronchoscope camera while the VS and Bronchoscope feeds are active in Dual View mode. The calculated average latency shall be less than 180 ms. | Bronchoscope average latency in Dual View mode: _____ ms Is the average latency less than 180 ms? _____
360 | Install the custom "provu" latency-test binary from the "Latency Test Binary" attachment folder on the latest hardware. 1. Copy the binary to a USB drive and connect the drive to the device. 2. Connect to the device through the serial terminal. 3. Execute the following commands in sequence: $ rm -f /usr/bin/provu $ mount /dev/sda1 /mnt $ cp /mnt/provu /usr/bin/provu $ chmod +x /usr/bin/provu $ sync Verify that each command completes without error and that /usr/bin/provu is executable. | The custom "provu" binary shall be copied to /usr/bin/provu and shall have executable permission. All commands shall complete without errors. | Is the USB drive mounted successfully? _____ Is the "provu" binary copied to /usr/bin/provu? _____ Is executable permission applied to the binary? _____ Does each command complete without errors? _____
370 | Reboot the device. Generate camera connection and disconnection events as follows: 1. Connect the Bronchoscope camera to Port A for 10 seconds, then disconnect it. 2. Connect the VS camera to Port B for 10 seconds, then disconnect it. 3. Connect the VL camera to Port C for 10 seconds, then disconnect it. 4. Start screen recording. 5. Reconnect the Bronchoscope, VS, and VL cameras to Ports A, B, and C, respectively. Keep them connected for 10 seconds, then disconnect all three cameras. | The device shall reboot successfully. Each camera shall be detected on the correct port and removed without error. Screen recording shall start successfully, and all camera connection and disconnection events shall be captured by the latency-test software. | Does the device reboot successfully? _____ Is the Bronchoscope detected on Port A? _____ Is the VS camera detected on Port B? _____ Is the VL camera detected on Port C? _____ Does screen recording start successfully? _____ Does each camera connection and disconnection event complete without errors? _____
380 | Use the device serial terminal to navigate to the /tmp directory: $ cd /tmp Verify that the latency data file exists: $ ls -l /tmp/provu_frame_delta_ms Copy "provu_frame_delta_ms" to a USB drive for analysis. Analyze the recorded data for Ports A, B, and C, and calculate the average processing time, render queue time, and end-to-end latency for each port. | The file /tmp/provu_frame_delta_ms shall exist and contain latency records for Ports A, B, and C. The calculated averages for each applicable port shall meet these limits: Average processing time < 37.7 ms Render queue time < 29 ms End-to-end latency < 55 ms | Is /tmp/provu_frame_delta_ms created? _____ Is the file copied successfully for analysis? _____ Port A — Processing: _____ ms; Render queue: _____ ms; End-to-end: _____ ms Port B — Processing: _____ ms; Render queue: _____ ms; End-to-end: _____ ms Port C — Processing: _____ ms; Render queue: _____ ms; End-to-end: _____ ms Does each applicable average meet the specified limit? _____
Test ID: STS-ProvUVista-Integration-000030 Software item (under test): SRS-FUN-CAM-050
390 | Power ON the device with the custom software application that displays camera frame rate. Connect the Bronchoscope camera to Port A, VS camera to Port B, and VL camera to Port C. Login with admin user and navigate to the live camera display screen. Verify that the application can show the frame-rate indication for each active camera feed.( On Screen This should be right side of number ) Note: The software update package is added with name "update-bundle-provu" in attachment , update that software on device to measure the fps. | Device shall boot successfully and allow admin login. Device shall detect Bronchoscope, VS, and VL cameras. Device shall display two active camera feeds as per the selected configuration and show frame-rate indication for the active feeds. | Does the device boot and allow admin login? _____ Are Bronchoscope, VS, and VL cameras detected? _____ Are two active camera feeds displayed? _____ Is frame-rate indication visible for active feeds? _____
400 | Enable screen recording while Bronchoscope and VS camera feeds are active. Verify that recording starts and the displayed frame-rate indication remains available for both active feeds. | Device shall start screen recording without interrupting the active camera feeds. Frame-rate indication shall remain visible for the active Bronchoscope and VS feeds during recording. | Does screen recording start successfully? _____ Are Bronchoscope and VS feeds uninterrupted? _____ Is frame-rate indication still visible during recording? _____
410 | Measure and record the displayed frame rate for Bronchoscope and VS camera feeds for a continuous 10-minute period while screen recording is enabled. Record minimum and average FPS for each feed. | VS camera display frame rate shall be at least 15 fps. Bronchoscope camera display frame rate shall be at least 25 fps. No feed freeze, drop-out, or loss of frame-rate indication shall occur during the 10-minute observation. | VS minimum FPS: _____ VS average FPS: _____ Bronchoscope minimum FPS: _____ Bronchoscope average FPS: _____ Did VS remain >= 15 fps? _____ Did Bronchoscope remain >= 25 fps? _____ Any freeze/drop-out observed? _____
420 | Deactivate the VS camera feed and activate the VL camera feed while keeping the Bronchoscope camera active. Measure and record the displayed frame rate for Bronchoscope and VL camera feeds for a continuous 10-minute period. | VL camera display frame rate shall be at least 15 fps. Bronchoscope camera display frame rate shall be at least 25 fps. Switching from VS to VL shall not interrupt the Bronchoscope feed or remove frame-rate indication. | VL minimum FPS: _____ VL average FPS: _____ Bronchoscope minimum FPS: _____ Bronchoscope average FPS: _____ Did VL remain >= 15 fps? _____ Did Bronchoscope remain >= 25 fps? _____ Was Bronchoscope feed uninterrupted during switching? _____
430 | Verify frame rate with each camera type active alone: VS, VL, and Bronchoscope. Then, with screen recording ON, connect a faulty camera on Port B while Bronchoscope is active and verify Bronchoscope frame rate. Repeat by connecting a faulty camera on Port C while VS is active, then while VL is active, and verify the healthy camera frame rate and fault indication. | When active alone or with a faulty camera connected, healthy camera feeds shall continue to meet the frame-rate requirement: VS camera >= 15 fps. VL camera >= 15 fps. Bronchoscope camera >= 25 fps. Faulty camera connection shall be indicated without degrading the active healthy camera feed below the required frame rate. | Single VS FPS min/avg: _____ / _____ Pass >=15 fps? _____ Single VL FPS min/avg: _____ / _____ Pass >=15 fps? _____ Single Bronchoscope FPS min/avg: _____ / _____ Pass >=25 fps? _____ Faulty camera on Port B indicated? _____ Bronchoscope FPS min/avg: _____ / _____ Pass >=25 fps? _____ Faulty camera on Port C indicated with VS active? _____ VS FPS min/avg: _____ / _____ Pass >=15 fps? _____ Faulty camera on Port C indicated with VL active? _____ VL FPS min/avg: _____ / _____ Pass >=15 fps? _____
Test ID: STS-ProvUVista-Integration-000040 Software item (under test): SRS-FUN-CAM-060
440 | Power ON the device. Connect the Bronchoscope camera to Port A. Connect the Logic analyzer to I2C Port 5. | Device should be powered ON. Logic analyzer should display the communication along with the register values being set. | Is the device powered ON? _____ Does the logic analyzer display the communication with register values being set? _____
450 | Keep the room lights ON and check the below registers value from logic analyzer. 0x3500 0x3501 0x3502 | Logic analyzer should display the registers with value. 0x3500 0x3501 0x3502 | Does the logic analyzer display the registers with values? _____ Recorded register values (lights ON): 0x3500: _____ 0x3501: _____ 0x3502: _____
460 | Turn OFF the room lights and reconnect the bronchoscope and check the values of the registers. 0x3500 0x3501 0x3502 | Device should change the values of the given registers. 0x3500 0x3501 0x3502 Device shall automatically adjust its exposure to compensate for changes in ambient lighting without user intervention. | Did the register values change? _____ Recorded register values (lights OFF): 0x3500: _____ 0x3501: _____ 0x3502: _____ Did the device automatically adjust exposure without user intervention? _____
470 | Connect the VL camera to Port B and VS camera to Port C. Deselect the bronchoscope camera and select the VS and VL camera feed on the device. | Device should display the VS and VL camera feed on the display. | Are the VS and VL camera feeds displayed on the device? _____ Does Device display the VS and VL camera feed on the display.
480 | Execute the command: v4l2-ctl -d /dev/video9 --list-ctrls-menu v4l2-ctl -d /dev/video11 --list-ctrls-menu Verify the supported values for exposure_auto Note: Use the VS camera having exposer configurable settings. | Auto Exposure control shall be available. exposure_auto 0x009a0901 (menu) : min=0 max=3 default=3 value=3 (Aperture Priority Mode) 1: Manual Mode 3: Aperture Priority Mode | Is the Auto Exposure control available? _____ Recorded exposure_auto value: /dev/video9: _____ /dev/video11: _____ Does device display the same controls as below: exposure_auto 0x009a0901 (menu) : min=0 max=3 default=3 value=3 (Aperture Priority Mode)
490 | Execute the command: v4l2-ctl -d /dev/video9 --set-ctrl=exposure_auto=1 v4l2-ctl -d /dev/video11 --set-ctrl=exposure_auto=1 Verify the supported values for exposure Execute the command: v4l2-ctl -d /dev/video9 --list-ctrls-menu v4l2-ctl -d /dev/video11 --list-ctrls-menu | Camera shall switch to Manual Exposure mode. exposure_auto 0x009a0901 (menu) : min=0 max=3 default=3 value=1 (Manual Mode) | Was the command executed successfully? _____ Is the camera in Manual Exposure mode? Recorded exposure_auto value: /dev/video9: _____ /dev/video11: _____ Does device display the same controls as below: Camera shall switch to Manual Exposure mode. exposure_auto 0x009a0901 (menu) : min=0 max=3 default=3 value=1 (Manual Mode)
500 | Reduce the ambient light or partially cover the camera lens. | The image shall become darker and shall not automatically adjust to compensate for the reduced lighting. | Does the image remain dark without automatic adjustment? _____ Does device behave the same as below: The image shall become darker and shall not automatically adjust to compensate for the reduced lighting.
510 | Restore normal lighting or expose the camera to a bright scene. | The image brightness shall not automatically adjust. | Does the image remain without automatic exposure adjustment? _____ Does device behave the same as below: The image brightness shall not automatically adjust.
520 | Execute the command: v4l2-ctl -d /dev/video9 --set-ctrl=exposure_auto=3 v4l2-ctl -d /dev/video11 --set-ctrl=exposure_auto=3 Verify the supported values for exposure Execute the command: v4l2-ctl -d /dev/video9 --list-ctrls-menu v4l2-ctl -d /dev/video11 --list-ctrls-menu | Camera shall switch back to Auto Exposure mode. exposure_auto 0x009a0901 (menu) : min=0 max=3 default=3 value=3 (Aperture Priority Mode) | Was the command executed successfully? _____ Is the camera back in Auto Exposure mode? Recorded exposure_auto value: /dev/video9: _____ /dev/video11: _____ Does device display the same controls as below: Camera shall switch back to Auto Exposure mode. exposure_auto 0x009a0901 (menu) : min=0 max=3 default=3 value=3 (Aperture Priority Mode)
530 | Reduce the ambient light or partially cover the camera lens. | The image shall automatically adjust to compensate for the reduced lighting for both of the camera feeds. | Does the image adjust automatically for both camera feeds? _____ VS feed adjusts automatically? _____ VL feed adjusts automatically? _____ Does device behave the same as below: The image shall automatically adjust to compensate for the reduced lighting for both of the camera feeds.
Test ID: STS-ProvUVista-Integration-000050 Software item (under test): SRS-FUN-CAM-070
540 | Execute: cat /sys/class/graphics/fb0/virtual_size | Framebuffer resolution shall be 1920,1080. | Recorded virtual_size: _____ Is the framebuffer resolution 1920 × 1080? _____
550 | Execute: cat /sys/class/graphics/fb0/bits_per_pixel | Framebuffer shall report 32 bits per pixel or 24 bits per pixel, depending on platform implementation. | Recorded bits_per_pixel: _____ Is the framebuffer 32 bpp or 24 bpp? _____
560 | Execute: fbset | Geometry shall report 1920 × 1080. | Recorded geometry (fbset): _____ Is the geometry 1920 × 1080? _____
570 | Verify the fbset | RGB channels shall each be 8 bits (RGB888). | Recorded rgba (fbset): _____ Are all RGB channels 8 bits each (RGB888)? _____
580 | Verify that the application UI is displayed correctly across the entire screen without distortion, scaling, clipping, or colour artifacts. | Graphics shall be displayed at 1920 × 1080 using RGB888 colour representation. | Does the application display correctly? _____
Test ID: STS-ProvUVista-Integration-000060 Software item (under test): SRS-FUN-CAM-080
590 | Connect the camera (VS/VL/Bronchoscope). Activate two camera feed Bronchoscope and VS Place a calibration grid or ruler in front of the camera | Setup | Setup
600 | Select x1 zoom and record the visible field of view or capture a screenshot. Verify that the default scaling factor is x1. | The live camera image shall be displayed at x1 scaling by default. | Is the default scaling factor x1? _____
610 | Select x1.5 zoom and measure the same object or field of view. | The live image shall be displayed at 1.5× zoom. The object's displayed size shall increase by approximately 1.5×, And the visible field of view shall decrease accordingly. | Is the image displayed at x1.5? _____
620 | Verify that the zoom is centered on the image using a test target placed at the center of the camera field of view. | The image shall be magnified about the image center without shifting the center position. | Is the zoom centered? _____
630 | Verify that the image aspect ratio is preserved at x1.5 zoom using a calibration grid Note: calculate the width/hight for checking the spect ratio. it should be same for x1 and x1.5 and x2. | The image shall not appear stretched, compressed, or geometrically distorted. | Is the aspect ratio preserved? _____
640 | Select the x2 zoom option. Select x2 zoom and measure the same object or field of view. | The live image shall be displayed at 2× zoom. The object's displayed size shall increase by approximately 2×, or the visible field of view shall decrease accordingly. | Is the image displayed at x2? _____
650 | Verify that the zoom remains centered and the aspect ratio is preserved at x2 zoom. | The image shall remain centered with no geometric distortion. | Is the image centered and undistorted? _____
660 | Repeat Steps for the VS Camera. | All supported camera sources shall provide x1.5 and x2 scaling with preserved aspect ratio. | Does Vs camera types pass? _____
Test ID: STS-ProvUVista-Integration-000070 Software item (under test): SRS-FUN-SFT-200
670 | Power ON the device and wait until the application is fully launched. Connect the Bronchoscope camera to Port A, the VS camera to Port B, and the VL camera to Port C. | Device should boot successfully and detect all three connected cameras. | Does the device boot successfully? _____ Are Bronchoscope, VS, and VL cameras detected? _____
680 | Navigate to the live camera display screen and verify that two camera feed images are displayed as PORTA and PORT B. | Software shall display two camera images on the screen. | Are a two camera images displayed on the screen for Bronchoscope and VS? _____
690 | Place an orientation test card with clear TOP, BOTTOM, LEFT, and RIGHT markings in front of each camera. Observe the displayed images in normal device position. | Each camera image shall be displayed in fixed landscape orientation and in the same perspective as the user. | Are all three images displayed in landscape orientation? _____ Are all images shown in the same user perspective? _____
700 | Verify that none of the displayed camera images are reversed or shown as a mirror image by comparing the LEFT/RIGHT markings on the test card. Deactivate the VS camera and Activate the Vl camera and test. | Software shall not display any camera image as reversed or mirror-image. | Is Bronchoscope image not mirrored? _____ Is VS image not mirrored? _____ Is VL image not mirrored? _____
710 | Physically rotate the device 90° clockwise while keeping the cameras and orientation test cards fixed. Observe all displayed camera images. Perform the test for both VS and Vl camera by deactivating and activating other. Note: See the attached snap "device_rotation_test_orientations" for example rotation of the device. | All camera images shall remain fixed in landscape orientation and shall not rotate when the device is rotated. | After 90° clockwise rotation, do all images remain fixed in landscape orientation? _____ Did any image rotate? _____
720 | Physically rotate the device 90° counter-clockwise while keeping the cameras and orientation test cards fixed. Observe all displayed camera images. | All camera images shall remain fixed in landscape orientation and shall not rotate when the device is rotated. | After 90° counter-clockwise rotation, do all images remain fixed in landscape orientation? _____ Did any image rotate? _____
730 | Physically rotate the device 180° while keeping the cameras and orientation test cards fixed. Observe all displayed camera images. | All camera images shall remain fixed in landscape orientation and shall not rotate when the device is rotated. | After 180° rotation, do all images remain fixed in landscape orientation? _____ Did any image rotate? _____
740 | Return the device to normal position and record the final result for SRS-FUN-SFT-200 image orientation safety. | Test shall pass only if all three camera images remain fixed in landscape orientation and are displayed in the same user perspective without reversal or mirror-image behavior for all tested device orientations. | Final result for SRS-FUN-SFT-200 Image orientation – Safety: Pass / Fail _____ Comments / evidence reference: _____
Test ID: STS-ProvUVista-Integration-000080 Software item (under test): SRS-FUN-GUI-210
750 | Navigate to Settings → Display → LCD Brightness. | Setup | Setup
760 | Verify that the brightness control allows adjustment from 0 to 100 and that the default value is 50. | Brightness control shall support values from 0 to 100. Default value shall be 50 (50% of maximum brightness). | Is the range 0–100 and default value 50? _____
770 | Keep the LCD brightness at default 50% and Execute: v4l2-ctl -d /dev/video0 -C lcd_brightness | The reported value shall be approximately 128 (50% of the maximum value of 255). | LCD Brightness Value: _____ Does LCD brightness approximately 128 (50% of the maximum value of 255)?____
780 | Set the UI brightness to 100. Execute: v4l2-ctl -d /dev/video0 -C lcd_brightness | The LCD brightness shall be applied within 1 second. The reported value shall be 255 (or the maximum supported value). | Does the LCD brightness Value 255?__ Does the brightness updated within 1 second? _____
790 | Set the UI brightness to 0. Execute: v4l2-ctl -d /dev/video0 -C lcd_brightness | The LCD backlight shall turn OFF within 1 second. The reported value shall be 0. | Did the LCD turn OFF within 1 second? _____ Brightness Value: ____
Test ID: STS-ProvUVista-Integration-000090 Software item (under test): SRS-FUN-GUI-220
800 | Connect two supported cameras VL + Bronchoscope on Port B and Port A | Setup | Setup
810 | Verify that live video from both cameras is displayed simultaneously. | Live video from both cameras shall be displayed correctly. | Are both live feeds displayed? _____
820 | Start screen recording while both camera feeds remain active. | Screen recording shall start successfully without interrupting the live camera feeds. | Did recording start successfully? _____
830 | Place a stopwatch (running milliseconds) or rapidly moving object in front of one camera. Observe the display while recording is active. | The displayed image shall follow the movement continuously without displaying old or frozen frames beyond the defined latency. | Is the displayed image continuously updated without stale frames? _____
840 | Repeat Step 830 for the second camera while screen recording remains enabled. | The displayed image shall follow the movement continuously without displaying old or frozen frames beyond the defined latency. | Is the displayed image continuously updated without stale frames? _____
850 | Stop the moving object and observe the display. | The live display shall immediately reflect the current scene and shall not display previously captured frames beyond the specified latency. | Are only current frames displayed? _____
860 | Stop screen recording and verify that both live camera feeds continue to display normally. | Live camera feeds shall continue without interruption after recording stops. | Do both live feeds continue normally? _____
870 | Verify the Steps 810 to 830 for Vs camera on Port C | Device should perform the same function for Vs camera on Port C as per steps 810 to 830. | Does Device perform the same function for Vs camera on Port C as per steps 820 to 840.
Test ID: STS-ProvUVista-Integration-000100 Software item (under test): SRS-FUN-CAM-1460
880 | Connect two supported cameras VL + Bronchoscope on Port B and Port A | Setup | Setup
890 | Verify that the live image is displayed correctly without any camera-related error. | Live camera feed shall be displayed normally. | Is the live feed displayed correctly? _____
900 | Simulate frame loss by breaking the data wire of the bronchoscope camera connection. | Setup | Setup
910 | Observe the Device screen with error And Verify the log from error log page. | The software shall detect the frame loss and display a camera frame loss error. Device should display the Camrea Error screen on display. Device shall display the Error text as "Fault detected on camera Port 1. Check camera and connections." Device shall display the Error log ID: 00029 with below texts "Fault with camera port#1 ( Accessories )" | Is Camera Error screen displayed? _____ Displayed error text: "Fault detected on camera Port 1. Check camera and connections." _____ Error log ID recorded: _____ Is Error ID 00029 logged? _____ Error log text: "Fault with camera port#1 (Accessories)" _____
920 | Connect back the data wire of the bronchoscope. Observe the camera feed resumes. | The live camera feed shall resume automatically, and the frame loss error shall be cleared. | Did the live feed resume and the error clear automatically? _____
930 | Simulate frame loss by breaking the data wire of the VL camera connected on Port B. | The software shall detect the frame loss and display a camera frame loss error. Device should display the Camrea Error screen on display. Device shall display the Error text as "Fault detected on camera Port 2. Check camera and connections." Device shall display the Error log ID: 00030 with below texts "Fault with camera port#2 ( Accessories )" | Is Camera Error screen displayed? _____ Displayed error text: "Fault detected on camera Port 2. Check camera and connections." _____ Error log ID recorded: _____ Is Error ID 00030 logged? _____ Error log text: "Fault with camera port#2 (Accessories)" _____
940 | Connect back the data wire of the VL camera. Observe the camera feed resumes. | The live camera feed shall resume automatically, and the frame loss error shall be cleared. | Did the live feed resume and the error clear automatically? _____
950 | Connect the VS camera on Port C and activate the feed and Simulate frame loss by breaking the data wire of the Vs camera connected on Port C. | The software shall detect the frame loss and display a camera frame loss error. Device should display the Camrea Error screen on display. Device shall display the Error text as "Fault detected on camera Port 3. Check camera and connections." Device shall display the Error log ID: 00031 with below texts "Fault with camera port#3 ( Accessories )" | Is Camera Error screen displayed? _____ Displayed error text: "Fault detected on camera Port 3. Check camera and connections." _____ Error log ID recorded: _____ Is Error ID 00031 logged? _____ Error log text: "Fault with camera port#3 (Accessories)" _____
960 | Connect back the data wire of the VS camera. Observe the camera feed resumes. | The live camera feed shall resume automatically, and the frame loss error shall be cleared. | Did the live feed resume and the error clear automatically? _____
Test ID: STS-ProvUVista-Integration-000110 Software item (under test): SRS-FUN-CAM-1820 , SRS-FUN-CAM-1840 , SRS-FUN-GUI-270
970 | Power ON the device and launch the application. Ensure no cameras are connected. Execute: cat /sys/devices/virtual/gpio-detection/oha428-camera/status cat /sys/devices/virtual/gpio-detection/usb-camera1/status cat /sys/devices/virtual/gpio-detection/usb-camera0/status Ensure the device audio/notification sound is ENABLED in settings and the volume is audible. No connect/remove tone is expected while no camera activity occurs. | All GPIO status values shall be 1 (Disconnected). All camera selection buttons shall be disabled and no live image shall be displayed. Device audio shall be enabled. No camera attach/remove notification tone shall be produced while no camera is connected or removed. | Are all GPIO status values 1 (Disconnected)? _____ Recorded status: oha428-camera: _____ usb-camera1: _____ usb-camera0: _____ Are all camera selection buttons disabled? _____ Is no live image displayed? _____ Is device audio enabled? _____ Was no attach/remove tone produced while idle? _____
980 | Connect the Bronchoscope to Port A. check the device port status and Within 2 seconds execute: cat /sys/devices/virtual/gpio-detection/oha428-camera/status v4l2-ctl --list-device Verify audio feedback while connecting bronchoscope camera. With audio ENABLED, listen for the audible notification tone at the moment the Bronchoscope is attached. Note : Verify the device audio with playing the attached file in folder "Audio Camera connect" | GPIO status shall change from 1 → 0. The Bronchoscope shall be detected as a V4L2 device, the Port A camera selection button shall become active, and the live image shall be displayed within 2 seconds. Device shall sound the camera connection with audio feedback. When audio is enabled, the software shall generate an audible notification tone as the Bronchoscope is attached. Device shall active of PORTA on screen within 2 seconds. | Did GPIO status change from 1 → 0? _____ Recorded oha428-camera status: _____ Is the Bronchoscope detected as a V4L2 device? _____ Detected device node: _____ Did the Port A selection button become active? _____ Was the live image displayed within 2 seconds? _____ Was audio feedback heard on connection? _____ Was an audible notification tone heard on Bronchoscope attach? _____ Does device activate the PORTA on screen with in 2 seconds.
990 | Disconnect the Bronchoscope from Port A. Within 2 seconds execute: cat /sys/devices/virtual/gpio-detection/oha428-camera/status With audio ENABLED, listen for the audible notification tone at the moment the Bronchoscope is removed. Verify the PORT A status on Screen. Note : Verify the device audio with playing the attached file in folder "Audio Camera connect" | GPIO status shall change from 0 → 1. The Port A camera selection button shall be disabled and the display frame shall be cleared within 2 seconds. When audio is enabled, the software shall generate an audible notification tone as the Bronchoscope is removed. Device shall de-active of PORTA on screen with in 2 seconds. | Did GPIO status change from 0 → 1? _____ Recorded oha428-camera status: _____ Was the Port A selection button disabled? _____ Was the display frame cleared within 2 seconds? _____ Was an audible notification tone heard on Bronchoscope removal? _____ Does device de-activate the PORTA on screen with in 2 seconds.
1000 | Connect the VS Camera to Port B. Within 2 seconds execute: cat /sys/devices/virtual/gpio-detection/usb-camera1/status v4l2-ctl --list-device With audio ENABLED, listen for the audible notification tone at the moment the VS camera is attached. Verify the PORT B status on Screen. Note : Verify the device audio with playing the attached file in folder "Audio Camera connect" | GPIO status shall change from 1 → 0. The VS camera shall be detected as a V4L2 device, the Port B camera selection button shall become active, and the live image shall be displayed within 2 seconds. When audio is enabled, the software shall generate an audible notification tone as the VS camera is attached. Device shall active of PORTB on screen with in 2 seconds. | Did GPIO status change from 1 → 0? _____ Recorded usb-camera1 status: _____ Is the VS camera detected as a V4L2 device? _____ Detected device node: _____ Did the Port B selection button become active? _____ Was the live image displayed within 2 seconds? _____ Was an audible notification tone heard on VS camera attach? _____ Does device activate the PORTB on screen with in 2 seconds.
1010 | Disconnect the VS Camera from Port B. Within 2 seconds execute: cat /sys/devices/virtual/gpio-detection/usb-camera1/status With audio ENABLED, listen for the audible notification tone at the moment the VS camera is removed. Verify the PORT B status on Screen. Note : Verify the device audio with playing the attached file in folder "Audio Camera connect" | GPIO status shall change from 0 → 1. The Port B camera selection button shall be disabled and the display frame shall be cleared within 2 seconds. When audio is enabled, the software shall generate an audible notification tone as the VS camera is removed. Device shall de-active of PORTB on screen with in 2 seconds. | Did GPIO status change from 0 → 1? _____ Recorded usb-camera1 status: _____ Was the Port B selection button disabled? _____ Was the display frame cleared within 2 seconds? _____ Was an audible notification tone heard on VS camera removal? _____ Does device de-activate the PORTB on screen with in 2 seconds.
1020 | Connect the VL Camera to Port C. Within 2 seconds execute: cat /sys/devices/virtual/gpio-detection/usb-camera0/status v4l2-ctl --list-device With audio ENABLED, listen for the audible notification tone at the moment the VL camera is attached. Verify the PORT C status on Screen. Note : Verify the device audio with playing the attached file in folder "Audio Camera connect" | GPIO status shall change from 1 → 0. The VL camera shall be detected as a V4L2 device, the Port C camera selection button shall become active, and the live image shall be displayed within 2 seconds. When audio is enabled, the software shall generate an audible notification tone as the VL camera is attached. Device shall active of PORTC on screen with in 2 seconds. | Did GPIO status change from 1 → 0? _____ Recorded usb-camera0 status: _____ Is the VL camera detected as a V4L2 device? _____ Detected device node: _____ Did the Port C selection button become active? _____ Was the live image displayed within 2 seconds? _____ Was an audible notification tone heard on VL camera attach? _____ Does device activate the PORTC on screen with in 2 seconds.
1030 | Disconnect the VL Camera from Port C. Within 2 seconds execute: cat /sys/devices/virtual/gpio-detection/usb-camera0/status With audio ENABLED, listen for the audible notification tone at the moment the VL camera is removed. Verify the PORT B status on Screen. Note : Verify the device audio with playing the attached file in folder "Audio Camera connect" | GPIO status shall change from 0 → 1. The Port C camera selection button shall be disabled and the display frame shall be cleared within 2 seconds. When audio is enabled, the software shall generate an audible notification tone as the VL camera is removed. Device shall de-active of PORTC on screen with in 2 seconds. | Did GPIO status change from 0 → 1? _____ Recorded usb-camera0 status: _____ Was the Port C selection button disabled? _____ Was the display frame cleared within 2 seconds? _____ Was an audible notification tone heard on VL camera removal? _____ Does device de-activate the PORTC on screen with in 2 seconds.
1040 | Connect the Bronchoscope (Port A), VS Camera (Port B) and VL Camera (Port C) simultaneously. Execute: cat /sys/devices/virtual/gpio-detection/oha428-camera/status cat /sys/devices/virtual/gpio-detection/usb-camera1/status cat /sys/devices/virtual/gpio-detection/usb-camera0/status v4l2-ctl --list-device With audio ENABLED, listen for an audible notification tone as each camera (Bronchoscope, VS, VL) is attached. Note : Verify the device audio with playing the attached file in folder "Audio Camera connect" | All GPIO status values shall be 0 (Connected). All three cameras shall be detected, all corresponding camera selection buttons shall become active within 2 seconds, and each live camera feed shall be available for selection. When audio is enabled, the software shall generate an audible notification tone for each camera attach event (Bronchoscope, VS, VL). | Are all GPIO status values 0 (Connected)? _____ Recorded status: oha428-camera: _____ usb-camera1: _____ usb-camera0: _____ Are all three cameras detected? _____ Did all three selection buttons become active within 2 seconds? _____ Is each live camera feed available for selection? _____ Was an audible notification tone heard for each camera attach? Bronchoscope: _____ VS: _____ VL: _____
1050 | Negative Test: Disconnect each camera individually while the remaining cameras stay connected. After each disconnection execute the corresponding GPIO status command. With audio ENABLED, listen for an audible notification tone at each individual camera removal. Note : Verify the device audio with playing the attached file in folder "Audio Camera connect" | Only the disconnected camera GPIO shall change from 0 → 1. Only its camera selection button shall be disabled and its display cleared. The remaining connected cameras shall continue streaming normally. When audio is enabled, the software shall generate an audible notification tone for each individual camera removal event. | For each disconnection, did only the disconnected camera GPIO change from 0 → 1? _____ Was only its selection button disabled and its display cleared? _____ Did the remaining connected cameras continue streaming normally? _____ Was an audible notification tone heard on each individual camera removal? _____
1060 | Negative Test: Connect an unsupported USB camera to Port B, then repeat on Port C. Execute: cat /sys/devices/virtual/gpio-detection/usb-camera1/status (or usb-camera0/status) v4l2-ctl --list-device With audio ENABLED, observe whether any notification tone is produced when the unsupported USB camera is attached and removed. | The GPIO shall indicate Connected (0) because hardware is attached. If the camera is unsupported, the software shall not activate the camera selection button or display a live image. When audio is enabled, an audible notification tone may be generated on hardware attach/removal; however, no camera feed shall be activated for the unsupported camera. | Did the GPIO indicate Connected (0) with the unsupported camera attached? _____ Recorded status (Port B / usb-camera1): _____ Recorded status (Port C / usb-camera0): _____ Did the software correctly NOT activate the camera selection button? _____ Was no live image displayed? _____ Was a tone produced for the unsupported camera attach/removal? _____ Was the camera feed correctly NOT activated? _____
1070 | Negative Test: Rapidly disconnect and reconnect each supported camera (Bronchoscope, VS, and VL) at least five times. v4l2-ctl --list-device Now DISABLE the device audio/notification sound in settings. Repeat rapidly disconnecting and reconnecting each supported camera (Bronchoscope, VS, VL) at least five times and listen for any tone. | The GPIO status shall correctly toggle between 1 and 0 for every connect/disconnect event. The software shall update the camera selection button and display within 2 seconds without crashes, stale images, or incorrect camera status. When audio is DISABLED, the software shall NOT generate any audible notification tone for any camera attach or removal event. | Did the GPIO status correctly toggle between 1 and 0 for every connect/disconnect event? _____ Did the software update the selection button and display within 2 seconds each time? _____ Were there any crashes, stale images, or incorrect camera status? _____ With audio DISABLED, was any notification tone produced? _____ (Expected: no tone)
1080 | Negative Test: Rapidly disconnect and reconnect each supported camera (Bronchoscope, VS, and VL) several times. Re-ENABLE the device audio/notification sound and rapidly disconnect and reconnect each supported camera several times, listening for a tone on every attach and removal event. Note : Verify the device audio with playing the attached file in folder "Audio Camera connect" | The software shall consistently detect each connection and disconnection within 2 seconds, without crashing, hanging, or leaving an incorrect camera status or stale image on the display. When audio is re-enabled, the software shall consistently generate an audible notification tone on every camera attach and removal event, without missed or spurious tones. | Did the software consistently detect each connection and disconnection within 2 seconds? _____ Were there any crashes, hangs, incorrect camera status, or stale images on the display? _____ With audio re-enabled, was a tone heard on every attach and removal event? _____ Any missed or spurious tones observed? _____
Test ID: STS-ProvUVista-Integration-000120 Software item (under test): SRS-FUN-CAM-2170
1090 | Ensure no bronchoscope is connected before starting the camera-type compatibility test. | Setup | Is the device powered ON and application launched? _____ Is no bronchoscope connected before starting? _____
1100 | Connect a bronchoscope with an OV6946 camera. Execute: cat /sys/devices/virtual/gpio-detection/oha428-camera/status cat /sys/bus/iio/devices/iio:device1/in_voltage0_raw v4l2-ctl -d /dev/video0 --list-formats-ext Verify detection, ADC ID, V4L2 enumeration, supported format, and live feed. | The Bronchoscope shall be detected and the feed shall be displayed on the device. The ADC value shall correspond to a supported bronchoscope ID. The camera shall be enumerated as a V4L2 device, and the live image shall be displayed without warning or error. Device shall display the supported format as 400x400. | GPIO status recorded: _____ ADC count recorded: _____ Does ADC correspond to supported bronchoscope ID? _____ Is OV6946 enumerated as a V4L2 device? _____ Detected device node: _____ Is supported format 400x400 displayed? _____ Is live image displayed without warning/error? _____
1110 | Disconnect the OV6946 bronchoscope and connect a bronchoscope with an OCHTA10-KL1x camera. Execute: cat /sys/devices/virtual/gpio-detection/oha428-camera/status cat /sys/bus/iio/devices/iio:device1/in_voltage0_raw v4l2-ctl -d /dev/video0 --list-formats-ext Verify detection, ADC ID, V4L2 enumeration, supported format, and live feed. | The Bronchoscope shall be detected and the feed shall be displayed on the device. The ADC value shall correspond to a supported bronchoscope ID. The camera shall be enumerated as a V4L2 device, and the live image shall be displayed without warning or error. Device shall display the supported format as 400x400. | GPIO status recorded: _____ ADC count recorded: _____ Does ADC correspond to supported bronchoscope ID? _____ Is OCHTA10-KL1x enumerated as a V4L2 device? _____ Detected device node: _____ Is supported format 400x400 displayed? _____ Is live image displayed without warning/error? _____
1120 | Verify camera functionality for both camera types by selecting the bronchoscope view and observing the live image for each camera type without restarting or changing software configuration. | The software shall acquire and display the live image correctly for both OV6946 and OCHTA10-KL1x cameras without requiring software restart or configuration changes. | Does OV6946 live image display correctly? _____ Does OCHTA10-KL1x live image display correctly? _____ Was software restart required? _____ Was configuration change required? _____
Test ID: STS-ProvUVista-Integration-000130 Software item (under test): SRS-FUN-PWR-1890 , SRS-FUN-CAM-2080
1130 | Keep the device powered OFF. Connect the logic analyzer to the I2C interface between the RK3568 and FPGA. Start the logic analyzer capture, then connect the Bronchoscope to the device. | The logic analyzer shall start capturing without error and shall be ready to capture RK3568-to-FPGA I2C communication before device power ON. | Is the device powered OFF before setup? _____ Is the logic analyzer connected to the RK3568-to-FPGA I2C interface? _____ Did the logic analyzer start capture without error? _____ Is the Bronchoscope connected to the device? _____
1140 | Power ON the device while the logic analyzer is actively capturing the I2C communication. | Setup. The device shall begin booting while the logic analyzer continues to capture the I2C communication. | Was the device powered ON successfully? _____ Did the logic analyzer continue capturing during power ON? _____
1150 | Verify that the device sends the required RK3568-to-FPGA configuration commands within 20 ms after power ON. Review the logic analyzer trace for register addresses and configured values. | The device shall send the FPGA configuration commands within 20 ms. The logic analyzer shall record the register list and configured values sent from RK3568 to FPGA. The device shall be able to configure the OV428 ISP chip. | Were FPGA configuration commands observed on the logic analyzer? _____ Measured time to first/complete configuration command: _____ ms Is the configuration completed within 20 ms? _____ Register addresses recorded: _____ Configured values recorded: _____
1160 | Verify that the device starts the application and removes the Bronchoscope emergency-mode feed during boot. | The device shall show the boot progress bar and remove the Bronchoscope emergency feed as the application starts. | Is the boot progress bar displayed? _____ Is the Bronchoscope emergency-mode feed removed during boot? _____ Did the application start without error? _____
1170 | Verify that the device completes booting and displays the home screen with the Bronchoscope camera connected. | The device shall complete booting and display the home screen with the Bronchoscope camera connected and available. | Did the device complete booting? _____ Is the home screen displayed? _____ Is the Bronchoscope camera connected and available? _____ Any boot or camera error observed? _____
1180 | Connect the logic analyzer on I2c port between rk3568 and ISP chip. Read and decode the resisters value. Verify the register 0x3503 and its value is 0x00 for OV428 ISP registers AGC (Auto Gain Control) and Auto exposure control (AEC) Verify the registers 0x5004 for Auto white balance is 0x90 (disabled). | The device shall send the configuration commands to ISP chip. The logic analyzer shall record the register list and configured values sent from RK3568 to ISP The device shall be able to configure the OV428 ISP chip. The device shall set OV428 ISP registers 0x3503 to 0x00 for AGC (Auto Gain Control) and Auto exposure control (AEC) to auto adjust the gain and exposure. The device shall disable the Auto white balance adjustment for OV428 by setting register 0x5004 to 0x90 | Were configuration commands observed on the logic analyzer? _____ Register addresses recorded: _____ Configured values recorded: _____ Were OV428 ISP configuration registers 0x3503 and 0x5004 observed? _____ Are AGC and AEC register 0x3503 set for 0x00 ? _____ Auto white balance register/value recorded 0x5004? _____ Is Auto White Balance register 0x5004 having value 0x90 ? _____
Test ID: STS-ProvUVista-Integration-000140 Software item (under test):SRS-FUN-USB-1660 , SRS-FUN-USB-1670
1190 | Preconditions: Device is powered ON and application is running. No USB cameras are connected. Export the Power-Good GPIOs using: echo 146 > /sys/class/gpio/export echo 148 > /sys/class/gpio/export Verify the initial Power-Good status using: cat /sys/class/gpio/gpio146/value cat /sys/class/gpio/gpio148/value | The device shall be powered ON and the application shall be running. No USB cameras shall be connected. GPIO146 and GPIO148 shall be exported successfully. Initial Power-Good status shall be LOW for both USB camera ports: GPIO148 = 0 for USB Camera Port B. GPIO146 = 0 for USB Camera Port C. | Is the device powered ON and application running? _____ Are no USB cameras connected? _____ Were GPIO146 and GPIO148 exported successfully? _____ Recorded initial Power-Good values: GPIO148 / Port B: _____ GPIO146 / Port C: _____ Are both Power-Good GPIO values 0? _____
1200 | Connect the VS camera to USB Camera (Port B ). Verify the camera detection and Power-Good signal using: cat /sys/devices/virtual/gpio-detection/usb-camera1/status cat /sys/class/gpio/gpio148/value | The VS camera shall be detected as connected on USB Camera Port B. The Port B Power-Good GPIO148 shall change from 0 to 1 within 2 seconds. The application shall enable the VS camera feed/selection and no camera fault or power warning shall be displayed. | Is the VS camera connected to USB Camera Port B? _____ Recorded usb-camera1 status: _____ Does the status indicate connected? _____ Recorded GPIO148 / Port B value: _____ Did GPIO148 change from 0 to 1 within 2 seconds? _____ Is the VS camera feed/selection enabled? _____ Any camera fault or power warning displayed? _____
1210 | Disconnect the VS camera from USB Camera PortB. Verify the camera detection and Power-Good signal using: cat /sys/devices/virtual/gpio-detection/usb-camera1/status cat /sys/class/gpio/gpio148/value | The VS camera shall be detected as disconnected from USB Camera Port B. The Port B Power-Good GPIO148 shall return to 0 within 2 seconds. The VS camera feed shall stop and the Port B camera selection shall be disabled. | Is the VS camera disconnected from USB Camera Port B? _____ Recorded usb-camera1 status: _____ Does the status indicate disconnected? _____ Recorded GPIO148 / Port B value: _____ Did GPIO148 return to 0 within 2 seconds? _____ Did the VS camera feed stop? _____ Is the Port B camera selection disabled? _____
1220 | Connect the VL camera to USB Camera (Port C ). Verify the camera detection and Power-Good signal using: cat /sys/devices/virtual/gpio-detection/usb-camera0/status cat /sys/class/gpio/gpio146/value | The VL camera shall be detected as connected on USB Camera Port C. The Port C Power-Good GPIO146 shall change from 0 to 1 within 2 seconds. The application shall enable the VL camera feed/selection and no camera fault or power warning shall be displayed. | Is the VL camera connected to USB Camera Port C? _____ Recorded usb-camera0 status: _____ Does the status indicate connected? _____ Recorded GPIO146 / Port C value: _____ Did GPIO146 change from 0 to 1 within 2 seconds? _____ Is the VL camera feed/selection enabled? _____ Any camera fault or power warning displayed? _____
1230 | Disconnect the VL camera from USB Camera PortC. Verify the camera detection and Power-Good signal using: cat /sys/devices/virtual/gpio-detection/usb-camera0/status cat /sys/class/gpio/gpio146/value | The VL camera shall be detected as disconnected from USB Camera Port C. The Port C Power-Good GPIO146 shall return to 0 within 2 seconds. The VL camera feed shall stop and the Port C camera selection shall be disabled. | Is the VL camera disconnected from USB Camera Port C? _____ Recorded usb-camera0 status: _____ Does the status indicate disconnected? _____ Recorded GPIO146 / Port C value: _____ Did GPIO146 return to 0 within 2 seconds? _____ Did the VL camera feed stop? _____ Is the Port C camera selection disabled? _____
1240 | Negative Test – USB Camera PortB: Connect the VS camera to USB Camera PortB. Introduce a hardware fault by breaking the signal U79_V2OK_GPIO4_C2 so that the USB port remains enabled but the Power-Good signal remains LOW. Verify using: cat /sys/class/gpio/gpio148/value | GPIO148 shall remain 0 while the VS camera is connected to USB Camera Port B. The software shall display the camera fault screen/text: "Fault detected on camera Port 2. Check camera and connections." Port 2 corresponds to USB Camera Port B. The fault shall be logged as Error ID 00035. | Recorded GPIO148 value: _____ Is GPIO148 still LOW while the Port B camera is connected? _____ Is Camera Error screen displayed? _____ Displayed fault text: "Fault detected on camera Port 2. Check camera and connections." _____ Is Port 2 confirmed as USB Camera Port B? _____ Error log ID recorded: _____ Is Error ID 00035 logged? _____
1250 | Negative Test – USB Camera PortC: Connect the VL camera to USB Camera PortC. Introduce a hardware fault by breaking the signal U78_V2OK_GPIO4_C2so that the USB port remains enabled but the Power-Good signal remains LOW. Verify using: cat /sys/class/gpio/gpio146/value | GPIO146 shall remain 0 while the VL camera is connected to USB Camera Port C. The software shall display the camera fault screen/text: "Fault detected on camera Port 3. Check camera and connections." Port 3 corresponds to USB Camera Port C. The fault shall be logged as Error ID 00036. | Recorded GPIO146 value: _____ Is GPIO146 still LOW while the Port C camera is connected? _____ Is Camera Error screen displayed? _____ Displayed fault text: "Fault detected on camera Port 3. Check camera and connections." _____ Is Port 3 confirmed as USB Camera Port C? _____ Error log ID recorded: _____ Is Error ID 00036 logged? _____
Test ID: STS-ProvUVista-Integration-000150 Software item (under test): SRS-FUN-CAM-1680-01 , SRS-FUN-CAM-1680-07 ,SRS-FUN-CAM-1690 ,SRS-FUN-CAM-1710
1260 | Connect the bronchoscope. Verify that the DS2482S-100 is responding. Execute: /usr/sbin/i2ctransfer -f -y 4 w2@0x1a 0x30 0x12 r1 | The bronchoscope shall be detected by the application. The terminal shall return 0x01. The response shall confirm communication with the DS2482S-100. | Was the bronchoscope detected by the application? _____ Recorded DS2482S-100 response: _____ Did the terminal return 0x01? _____ Is communication with the DS2482S-100 confirmed? _____
1270 | Verify that the EEPROM device is present. Execute: ls /sys/devices/w1_bus_master1/ | The 1-Wire EEPROM device shall be detected and its device ID shall be displayed in the expected format, for example "2d-000041e171f6". | Was the 1-Wire EEPROM device detected? _____ Recorded EEPROM device ID: _____ Is the device ID in the expected format? _____
1280 | Read the EEPROM contents. Execute: hexdump -v -C /sys/devices/w1_bus_master1/<deviceID>/eeprom Replace <deviceID> with the EEPROM device ID identified in Step 1270. | The EEPROM contents shall be displayed in hexadecimal format. | Were the EEPROM contents displayed in hexadecimal format? _____ EEPROM output / evidence reference: _____
1290 | Verify the EEPROM data format (<Header><Data><CRC>) against the EEPROM specification. Byte 0: EEPROM table version ID(Chip ID) Byte 1: Header Length Byte 5: CRC-1 (MSB) Byte 6: CRC-2 Byte 7: CRC-3 Byte 8: CRC-4 (LSB) Read the EEPROM data starting from byte 0. | The EEPROM data shall match the specified format: EEPROM table version ID = 0x01 Header Length = 0x09 (9 bytes) CRC shall be stored in bytes 5–8, with byte 5 as the MSB and byte 8 as the LSB. | Recorded EEPROM table version ID: _____ Recorded Header Length: _____ Recorded CRC bytes 5–8: _____ Does the EEPROM data match the specified format? _____
1300 | Calculate and verify the 32-bit CRC. Copy the attached script crc_eeprom.sh to the device and execute: hexdump -v -C /sys/devices/w1_bus_master1/<deviceID>/eeprom > eeprom_data.txt ./crc_eeprom.sh eeprom_data.txt Example output: CRC32 : 0x9ED09981 Compare the calculated CRC with the stored CRC in bytes 5–8. | The calculated 32-bit CRC shall match the CRC value stored in EEPROM bytes 5–8. | Calculated CRC32: _____ Stored CRC from bytes 5–8: _____ Do the calculated and stored CRC values match? _____
1310 | Verify detection of corrupted EEPROM data. 1. Create invalid data: printf '\x05\x06\x07\x08' > /tmp/eeprom.bin 2. Identify EEPROM device ID: ls /sys/devices/w1_bus_master1/ 3. Write invalid data: cat /tmp/eeprom.bin > /sys/devices/w1_bus_master1/<deviceID>/eeprom 4. Disconnect and reconnect the bronchoscope. 5. Read EEPROM: hexdump -v -C /sys/devices/w1_bus_master1/<deviceID>/eeprom 6. Verify the first four bytes and the application warning/error log. | The EEPROM shall be detected, but its contents shall fail validation. The application shall display: Faulty camera connected. Check device and connections. Error ID 00305 shall be logged for corrupted EEPROM data. Error ID 00306 shall be logged for an invalid EEPROM data format, where applicable. | Was the invalid EEPROM detected? _____ Recorded first four EEPROM bytes: _____ Was the warning displayed? _____ Displayed warning text: _____ Error ID recorded: _____ Is Error ID 00305 logged? _____ Is Error ID 00306 logged? _____
1320 | Verify operation with a valid EEPROM. Connect a bronchoscope containing a valid EEPROM. | The EEPROM shall be accepted successfully. The application shall not display the faulty-camera warning. Error ID 00305 shall not be logged. | Was the valid EEPROM accepted? _____ Was any faulty-camera warning displayed? _____ Error ID recorded: _____ Is Error ID 00305 absent? _____
1330 | Verify the default bronchoscope OV428 configuration using an I2C sniffer. Connect the I2C sniffer to I2C Port 5 and capture the bronchoscope configuration values during initialization. Re-connect the Bronchoscope while rung the logic analyzer software. Capture the I2c register value in Actual Received Value of "Default Configuration Broncho" sheet Compare the captured register values column with the values listed in the Default Configuration Broncho sheet. ( Expected Read vs Actual Received Value) | The device shall transmit the bronchoscope configuration values on I2C Port 5. The captured register values match the expected values in the Default Configuration Bronco sheet. | Were the configuration values captured on I2C Port 5? _____ Captured register values / evidence reference: _____ Do the values match the Default Configuration Bronco sheet? _____ Were the default OV428 settings used? _____
1340 | Verify behavior when the bronchoscope has no EEPROM. Connect a bronchoscope without an EEPROM device. Execute: ls /sys/devices/w1_bus_master1/ If an EEPROM device is reported, attempt: hexdump -v -C /sys/devices/w1_bus_master1/<deviceID>/eeprom Verify the application warning and error log. | The EEPROM shall not be available for reading. If the read command is executed against the unavailable EEPROM path, the terminal is expected to report: hexdump: all input file arguments failed The application shall display: Faulty camera connected. Check device and connections. Error ID 00306 shall be logged. | Was the EEPROM absent/not detected? _____ Recorded terminal error message: _____ Did the command report 'hexdump: all input file arguments failed'? _____ Was the faulty-camera warning displayed? _____ Displayed warning text: _____ Error ID recorded: _____ Is Error ID 00306 logged? _____
1350 | Re-verify the OV428 default configuration after EEPROM-related tests. Connect the I2C sniffer to I2C Port 5 and capture the bronchoscope configuration values. Compare the captured register values with the Default Configuration Bronco sheet. | The device shall transmit the expected bronchoscope configuration values on I2C Port 5 and shall use the default OV428 configuration. | Were configuration values captured on I2C Port 5? _____ Captured register values / evidence reference: _____ Do the values match the Default Configuration Bronco sheet? _____ Were the default OV428 settings used? _____
Test ID: STS-ProvUVista-Integration-000160 Software item (under test): SRS-FUN-CAM-1700, SRS-FUN-CAM-2030
1360 | 1. Connect the bronchoscope that contains the configured OV428 data in its EEPROM. 2. Identify the EEPROM device ID by executing: ls /sys/devices/w1_bus_master1/ 3. Read the EEPROM contents by executing: hexdump -v -C /sys/devices/w1_bus_master1/<deviceID>/eeprom 4. Save the EEPROM output as test evidence. 5. Locate the byte sequence "30 0e" and record the byte immediately following it as the configured value for OV428 register 0x300E. Example: In the sequence "30 0e e7", the configured register value is 0xE7. Note: The EEPROM device ID is displayed in a format such as "2d-000039bf4972". | The EEPROM device ID shall be displayed. The EEPROM contents shall be read successfully in hexadecimal format. The byte sequence "30 0e" and its following data byte shall be present in the EEPROM output. The configured value for OV428 register 0x300E shall be recorded. | Recorded EEPROM device ID: _____ Was the EEPROM read completed successfully? _____ Was the EEPROM output displayed in hexadecimal format? _____ Was the byte sequence "30 0e" found? _____ Configured value for OV428 register 0x300E: _____ Evidence/output reference: _____
1370 | 1. From the device serial terminal, read OV428 register 0x300E on I2C bus 5 by executing: /usr/sbin/i2ctransfer -y 5 w2@0x36 0x30 0x0e r1 2. Verify that the command completes successfully and returns one byte in hexadecimal format. 3. Record the returned register value. 4. Compare the returned value with the configured value recorded from the EEPROM in Step 1360. | The I2C read command shall complete successfully and return one hexadecimal byte from OV428 register 0x300E. The value read from register 0x300E shall match the configured EEPROM value recorded in Step 1360. | Did the I2C read command complete successfully? _____ Value read from OV428 register 0x300E: _____ Configured EEPROM value from Step 1360: _____ Do the two values match? _____ Evidence/terminal-output reference: _____
1380 | 1. Create the EEPROM binary file by executing: printf '\x01\x09\x26\x00\x00\x9e\xd0\x99\x81\x01\x14\x1a\x05\x15\x00\x00\x00\x00\x00\x00\x01\x00\x00\x00\x02\x01\x00\x00\x00\x00\x00\x00\x30\x0e\xe1\x36\x01\x00' > /tmp/eeprom.bin 2. Verify that /tmp/eeprom.bin was created successfully. 3. Identify the EEPROM device ID by executing: ls /sys/devices/w1_bus_master1/ 4. Write the binary file to the EEPROM by executing: cat /tmp/eeprom.bin > /sys/devices/w1_bus_master1/<deviceID>/eeprom 5. Verify that the write command completes without an error. Note: The written EEPROM configuration specifies the value 0xE1 for OV428 register 0x300E. | The EEPROM binary file shall be created successfully. The EEPROM device ID shall be displayed. The binary file shall be written to the selected EEPROM without errors. The written configuration shall specify 0xE1 as the value for OV428 register 0x300E. | Was /tmp/eeprom.bin created successfully? _____ Recorded EEPROM device ID: _____ Did the EEPROM write command complete without errors? _____ Configured value written for OV428 register 0x300E: _____ Evidence/output reference: _____
1390 | 1. Identify the EEPROM device ID by executing: ls /sys/devices/w1_bus_master1/ 2. Read the EEPROM contents by executing: hexdump -v -C /sys/devices/w1_bus_master1/<deviceID>/eeprom 3. Locate the byte sequence "30 0e". 4. Verify that the byte immediately following this sequence is "e1". Example: "30 0e e1" indicates that the EEPROM configuration value for OV428 register 0x300E is 0xE1. | The EEPROM device ID shall be displayed. The EEPROM contents shall be read successfully in hexadecimal format. The EEPROM output shall contain the sequence "30 0e e1", confirming that the configured value for OV428 register 0x300E is 0xE1. | Recorded EEPROM device ID: _____ Was the EEPROM read completed successfully? _____ Was the sequence "30 0e e1" found? _____ Evidence/output reference: _____
1400 | 1. Ensure the bronchoscope is connected and initialized using the EEPROM configuration verified in Step 1400. 2. From the device serial terminal, read OV428 register 0x300E on I2C bus 5 by executing: /usr/sbin/i2ctransfer -y 5 w2@0x36 0x30 0x0e r1 3. Verify that the command completes successfully and returns one byte in hexadecimal format. 4. Record the returned value and compare it with the EEPROM value verified in Step 1390. | The I2C read command shall complete successfully and return one hexadecimal byte from OV428 register 0x300E. The value read from register 0x300E shall be 0xE1 and shall match the EEPROM value verified in Step 1390. | Did the I2C read command complete successfully? _____ Value read from OV428 register 0x300E: _____ EEPROM value from Step 1400: _____ Is the register value 0xE1? _____ Do the register and EEPROM values match? _____ Evidence/terminal-output reference: _____
1410 | 1. Re-create the EEPROM binary file with the OV428 register 0x300E value set to 0xE7 by executing: printf '\x01\x09\x26\x00\x00\x9e\xd0\x99\x81\x01\x14\x1a\x05\x15\x00\x00\x00\x00\x00\x00\x01\x00\x00\x00\x02\x01\x00\x00\x00\x00\x00\x00\x30\x0e\xe7\x36\x01\x00' > /tmp/eeprom.bin 2. Verify that /tmp/eeprom.bin was created successfully. 3. Identify the EEPROM device ID by executing: ls /sys/devices/w1_bus_master1/ 4. Write the binary file to the EEPROM by executing: cat /tmp/eeprom.bin > /sys/devices/w1_bus_master1/<deviceID>/eeprom 5. Verify that the write command completes without an error. | The EEPROM binary file shall be created successfully. The EEPROM device ID shall be displayed. The binary file shall be written to the selected EEPROM without errors. The written configuration shall specify 0xE7 as the value for OV428 register 0x300E. | Was /tmp/eeprom.bin created successfully? _____ Recorded EEPROM device ID: _____ Did the EEPROM write command complete without errors? _____ Configured value written for OV428 register 0x300E: _____ Evidence/output reference: _____
1420 | 1. Identify the EEPROM device ID by executing: ls /sys/devices/w1_bus_master1/ 2. Read the EEPROM contents by executing: hexdump -v -C /sys/devices/w1_bus_master1/<deviceID>/eeprom 3. Locate the byte sequence "30 0e". 4. Verify that the byte immediately following this sequence is "e7". Example: "30 0e e7" indicates that the EEPROM configuration value for OV428 register 0x300E is 0xE7. | The EEPROM device ID shall be displayed. The EEPROM contents shall be read successfully in hexadecimal format. The EEPROM output shall contain the sequence "30 0e e7", confirming that the configured value for OV428 register 0x300E is 0xE7. | Recorded EEPROM device ID: _____ Was the EEPROM read completed successfully? _____ Was the sequence "30 0e e7" found? _____ Configured EEPROM value for OV428 register 0x300E: _____ Is the recorded value 0xE7? _____ Evidence/output reference: _____
1430 | 1. Ensure the bronchoscope is connected and initialized using the EEPROM configuration verified in Step 1430. 2. From the device serial terminal, read OV428 register 0x300E on I2C bus 5 by executing: /usr/sbin/i2ctransfer -y 5 w2@0x36 0x30 0x0e r1 3. Verify that the command completes successfully and returns one byte in hexadecimal format. 4. Record the returned value and compare it with the EEPROM value verified in Step 1420. | The I2C read command shall complete successfully and return one hexadecimal byte from OV428 register 0x300E. The value read from register 0x300E shall be 0xE7 and shall match the EEPROM value verified in Step 1420. | Did the I2C read command complete successfully? _____ Value read from OV428 register 0x300E: _____ EEPROM value from Step 1420: _____ Is the register value 0xE7? _____ Do the register and EEPROM values match? _____ Evidence/terminal-output reference: _____

## Latency Test Record
Stopwatch method calculator helper
Stopwatch Time on snap | Millisecs | Stopwatch Time on Display on same snap | Millisecs | Diff
00:00:120 | 120 | 00:00:20 | 20 | 100
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
| #VALUE! |  | #VALUE!
|  |  | Average | 100
Note: While entering the number insure the formatting should be same and remove the first entry.

## Aspect ratio Test
Table 1
Aspect ratio test for Port B
S. No. | Test Method | Expected Result | Actual Result
Table1-10 | Keep the camera connected to Port B. Verify that the camera feed is displayed in the center of the screen. | The device displays the camera feed and displays Port B as selected. | Does the device display the camera feed and display Port B as selected? ______
Table1-20 | Measure the width and height of the displayed camera feed using a scale. Record the measured values. Calculate the aspect ratio using Width ÷ Height and record the calculated value. | The width and height of the displayed camera feed are measured and recorded. The calculated aspect ratio is recorded. | Is the measured width = ______ and height = ______? Is the calculated aspect ratio Width ÷ Height = ______?
Table1-30 | Press the Zoom button twice to set the zoom level to 200%. Measure the width and height of the displayed camera feed using a scale. Record the measured values. Calculate the aspect ratio using Width ÷ Height and record the calculated value. | The camera feed is displayed at 200% zoom. The width and height of the displayed camera feed are measured and recorded. The calculated aspect ratio is recorded. | Does the camera feed change to 200% zoom? Is the measured width = ______ and height = ______? Is the calculated aspect ratio Width ÷ Height = ______?
Table1-40 | Verify that the aspect ratio calculated before zooming and after applying 200% zoom is the same. | The device displays the same aspect ratio before and after applying 200% zoom. | Is the aspect ratio before zoom ______ and after 200% zoom ______? Is the aspect ratio the same before and after zoom? ______
Table1-50 | Connect the camera to Port C to switch the device to Dual View mode. Measure the width and height of the Port B camera feed using a scale. Record the measured values. Calculate the aspect ratio using Width ÷ Height and record the calculated value. | The device displays the same aspect ratio for the Port B camera feed in Single View and Dual View modes. | Is the measured Port B width = ______ and height = ______? Is the calculated aspect ratio Width ÷ Height = ______? Is the aspect ratio the same as in Single View? ______
Table 2
Aspect ratio test for Port C
S. No. | Test Method | Expected Result | Actual Result
Table2-10 | Keep the camera connected to Port C. Verify that the camera feed is displayed in the center of the screen. | The device displays the camera feed and displays Port C as selected. | Does the device display the camera feed and display Port C as selected? ______
Table2-20 | Measure the width and height of the displayed camera feed using a scale. Record the measured values. Calculate the aspect ratio using Width ÷ Height and record the calculated value. | The width and height of the displayed camera feed are measured and recorded. The calculated aspect ratio is recorded. | Is the measured width = ______ and height = ______? Is the calculated aspect ratio Width ÷ Height = ______?
Table2-30 | Press the Zoom button twice to set the zoom level to 200%. Measure the width and height of the displayed camera feed using a scale. Record the measured values. Calculate the aspect ratio using Width ÷ Height and record the calculated value. | The camera feed is displayed at 200% zoom. The width and height of the displayed camera feed are measured and recorded. The calculated aspect ratio is recorded. | Does the camera feed change to 200% zoom? Is the measured width = ______ and height = ______? Is the calculated aspect ratio Width ÷ Height = ______?
Table2-40 | Verify that the aspect ratio calculated before zooming and after applying 200% zoom is the same. | The device displays the same aspect ratio before and after applying 200% zoom. | Is the aspect ratio before zoom ______ and after 200% zoom ______? Is the aspect ratio the same before and after zoom? ______
Table2-50 | Connect the camera to Port B to switch the device to Dual View mode. Measure the width and height of the Port C camera feed using a scale. Record the measured values. Calculate the aspect ratio using Width ÷ Height and record the calculated value. | The device displays the same aspect ratio for the Port C camera feed in Single View and Dual View modes. | Is the measured Port C width = ______ and height = ______? Is the calculated aspect ratio Width ÷ Height = ______? Is the aspect ratio the same as in Single View? ______
Table 3
Aspect ratio test for Port C
S. No. | Test Method | Expected Result | Actual Result
Table3-10 | Keep the camera connected to Port A. Verify that the camera feed is displayed in the center of the screen. | The device displays the camera feed and displays Port A as selected. | Does the device display the camera feed and display Port A as selected? ______
Table3-20 | Measure the width and height of the displayed camera feed using a scale. Record the measured values. Calculate the aspect ratio using Width ÷ Height and record the calculated value. | The width and height of the displayed camera feed are measured and recorded. The calculated aspect ratio is recorded. | Is the measured width = ______ and height = ______? Is the calculated aspect ratio Width ÷ Height = ______?
Table3-30 | Press the Zoom button twice to set the zoom level to 200%. Measure the width and height of the displayed camera feed using a scale. Record the measured values. Calculate the aspect ratio using Width ÷ Height and record the calculated value. | The camera feed is displayed at 200% zoom. The width and height of the displayed camera feed are measured and recorded. The calculated aspect ratio is recorded. | Does the camera feed change to 200% zoom? Is the measured width = ______ and height = ______? Is the calculated aspect ratio Width ÷ Height = ______?
Table3-40 | Verify that the aspect ratio calculated before zooming and after applying 200% zoom is the same. | The device displays the same aspect ratio before and after applying 200% zoom. | Is the aspect ratio before zoom ______ and after 200% zoom ______? Is the aspect ratio the same before and after zoom? ______
Table3-50 | Connect the camera to Port B to switch the device to Dual View mode. Measure the width and height of the Port A camera feed using a scale. Record the measured values. Calculate the aspect ratio using Width ÷ Height and record the calculated value. | The device displays the same aspect ratio for the Port A camera feed in Single View and Dual View modes. | Is the measured Port A width = ______ and height = ______? Is the calculated aspect ratio Width ÷ Height = ______? Is the aspect ratio the same as in Single View? ______

## Default Configuration Broncho
OV428 Register Readback Review
Register Address | Expected Read Value | Review Comment | Actual Received Value
0x0301 | 0x88
0x0304 | 0x00
0x0305 | 0xF0
0x0306 | 0x04
0x0307 | 0x04
0x0324 | 0x00
0x0325 | 0xF0
0x032A | 0x0E
0x032B | 0x02
0x3001 | 0x16
0x300D | 0x01
0x300E | Based On EEPROM
0x300F | 0x01
0x3013 | 0x31
0x3014 | 0x02
0x3015 | 0x00
0x3018 | 0xF0
0x301A | 0xAA
0x3031 | 0x02
0x3107 | 0x01
0x3200 | 0x00
0x3201 | 0x20
0x3202 | 0x00
0x3204 | 0xCE
0x3207 | 0x00
0x3208 | 0x10
0x3209 | 0x03
0x3215 | 0x03
0x3416 | 0x02
0x3418 | 0x00
0x3500 | 0x00 | Not applicable: this register is effective only in manual exposure mode; the current configuration uses auto-exposure mode.
0x3501 | 0x01 | Not applicable: this register is effective only in manual exposure mode; the current configuration uses auto-exposure mode.
0x3502 | 0xBE | Not applicable: this register is effective only in manual exposure mode; the current configuration uses auto-exposure mode.
0x3503 | 0x00
0x3508 | 0x03 | Not applicable: this register is effective only in manual gain mode; the current configuration uses auto-gain mode.
0x3509 | 0x00 | Not applicable: this register is effective only in manual gain mode; the current configuration uses auto-gain mode.
0x350A | 0x02 | Accepted exception: the OmniVision OVD file sets this register to 0x01, but the readback value is 0x02. The datasheet does not provide a clear definition.
0x350B | 0xD0 | Accepted exception: the OmniVision OVD file sets this register to 0x00, but the readback value is 0xD0. The datasheet does not provide a clear definition.
0x350C | 0x00 | Accepted exception: the OmniVision OVD file sets this register to 0x00 and the readback value is 0x00. The datasheet does not provide a clear definition.
0x3523 | 0x03
0x3524 | 0x00
0x3600 | 0x00
0x3601 | Based On EEPROM
0x3602 | 0x08
0x3603 | 0x50
0x3604 | 0x14
0x3605 | 0x40
0x3606 | 0x08
0x3607 | 0x00
0x3608 | 0x30
0x360E | 0x02
0x360F | 0x00
0x3611 | 0x40
0x3612 | 0x80
0x3617 | 0x08
0x3631 | 0x91
0x3660 | 0x80
0x3663 | 0x00
0x3664 | 0x00
0x3668 | 0x04
0x366A | 0xB5
0x3672 | 0x36
0x3674 | 0xF6
0x3701 | 0x70
0x3702 | 0x49
0x3708 | 0x0B
0x3709 | 0x11
0x370A | 0x0E
0x370B | 0x14
0x370C | 0x06
0x3730 | 0x11
0x3731 | 0x11
0x3732 | 0x11
0x3733 | 0xFF
0x3736 | 0x00
0x3737 | 0x88
0x373A | 0x01
0x3760 | 0x00 | Accepted exception: the OmniVision OVD file sets this register to 0x80, but the readback value is 0x00. The datasheet does not define this register.
0x3800 | 0x01
0x3805 | 0x01
0x3806 | 0xC4
0x3809 | 0x04
0x380B | 0x6C
0x3812 | 0x02
0x3813 | 0xDC
0x3820 | 0x00
0x3821 | 0x00
0x3822 | 0x00
0x3823 | 0x00
0x3824 | 0x02
0x3825 | 0xD7
0x3826 | 0x02
0x3827 | 0xD7
0x3828 | 0x01
0x3829 | 0x90
0x382A | 0x01
0x382B | 0x90
0x382C | 0x03
0x382D | 0x08
0x382E | 0x00
0x382F | 0x00
0x3831 | 0x00
0x3833 | 0x00
0x3840 | 0x80
0x3852 | 0x06
0x3853 | 0xBF
0x3854 | 0x02
0x3855 | 0xDC
0x3856 | 0x0A
0x3A00 | 0x18
0x3A02 | 0x01
0x3A03 | 0x90
0x3A05 | Based On EEPROM
0x3A0F | Based On EEPROM
0x3A10 | Based On EEPROM
0x3A11 | 0xD0
0x3A13 | 0x94
0x3A14 | Based On EEPROM
0x3A15 | Based On EEPROM
0x3A18 | Based On EEPROM
0x3A19 | Based On EEPROM
0x3A1A | 0x06
0x3A1B | Based On EEPROM
0x3A1E | Based On EEPROM
0x3D85 | 0x0B
0x3F05 | 0x00
0x4002 | 0x00
0x4003 | 0x60
0x4004 | 0x00
0x4005 | 0x02
0x4006 | 0x00
0x4007 | 0x02
0x4008 | 0x00
0x4009 | 0x01
0x400A | 0x03
0x400C | 0x00
0x400D | 0x01
0x4010 | 0xF0
0x4604 | 0x48
0x480E | 0x00
0x4813 | 0xE4
0x4827 | 0x00
0x4837 | 0x10
0x4B27 | 0x00
0x4F01 | 0x13
0x5004 | 0x90
0x5005 | 0x04
0x500C | 0x80
0x500D | 0x24
0x5010 | 0x01
0x5011 | 0x98
0x5012 | 0x01
0x5013 | 0x94
0x5016 | 0x01
0x5017 | 0x97
0x501A | 0x01
0x501B | 0x93
0x5020 | 0x01
0x5021 | 0x97
0x5022 | 0x01
0x5023 | 0x93
0x5044 | 0x04
0x5045 | 0x02
0x5048 | 0x01
0x5049 | 0x90
0x504A | 0x01
0x504B | 0x90
0x5050 | 0x00
0x5051 | 0x00
0x5052 | 0x00
0x5053 | 0x00
0x5054 | 0x01
0x5055 | 0x90
0x5056 | 0x01
0x5057 | 0x90
0x5058 | 0x01
0x5059 | 0x90
0x505A | 0x01
0x505B | 0x90