## CoverPage
ProVu Vista Display – Software Integration Test Request - 06
Project: |  | 155 - ProVu Accessories
Document ID/Version: |  | GTD-001228/A
DOCUMENT HISTORY
Version | Document Changes | Author | Date
A | Initial release | Ayush Pareek | 2026-09-22
Associated Documents
Document No | Document Title
SRS-000053 | ProVu Vista Display –SOFTWARE REQUIREMENT SPECIFICATION
GTD-000993 | PROVU2 LARGE DISPLAY Software Test Plan
Attachment
Name | Description and Summary
UpgradeBundle-NoEEPROMDriver | This is a test bundle, this does not have Bronchoscope EEPROM driver. Using this test bundle first we will verify that software does not detect the EEPROM
UpgradeBundle-ReleaseUSB | This is a release bundle. The release software contains the required EEPROM driver changes including kernel configuration.
UpgradeBundle-InvalidManifest-CorrectSha256 | This is a test bundle. This software package modifies the manifest contents, recalculates the SHA-256 checksum, and updates the checksum value
UpgradeBundle-OlderSoftwarePackage | This is a test bundle. This is having a earlier release version just to check, software detects the Older version and notify user.
UpgradeBundle-InvalidPCB-ID | This is a test bundle, provides invalid PCB ID configuration for testing.
UpgradeBundle-InvalidManifest | This is a test bundle, modifies the manifest contents but not update the sha256.
UpgradeBundle-IncorrectName | This is a test bundle, modifies the bundle name
UpgradeBundle-InvalidSizeLow | This is a test bundle, modifies the bundle contents to reduce the size.
UpgradeBundle-InvalidSizeHigh | This is a test bundle, modifies the bundle contents to increase the size.
UpgradeBundle-InvalidSignature | This is a test bundle, uses the different signature to sign the bundle.
Location | Description
Repository Name: GTD-001228 Folder Name: provu2_test_suite / GTD-001228 | https://dev.azure.com/FlexicareDevOps/ProVu%2015/_git/provu2_test_suite?version=Gbmain

## About
GTD-001228: Software Integration Test Request -06
Test Request Information
Author | Ayush Pareek
Revision | A
Reference | GTD-000993 | Date | 2026-09-22
Test objective
This test script verifies the software upgrade mechanism and file system protections.
Test Coverage
Software Upgrade: SRS-SEC-SUP-1030-01, SRS-SEC-SUP-1030-02, SRS-SEC-SUP-1030-03, SRS-SEC-SUP-1030-04, SRS-SEC-SUP-1030-05, SRS-SEC-SUP-1590, SRS-SEC-SUP-1610, SRS-SEC-SUP-1620, SRS-SEC-SUP-1800, SRS-SEC-SUP-1810, SRS-SEC-SUP-1870, SRS-SEC-SUP-1970, SRS-SEC-SUP-1980, SRS-SEC-SUP-2180, SRS-SEC-SUP-2190, SRS-SEC-SUP-1990, SRS-EXT-ITF-1490 Security, and privacy protection: SRS-SEC-HARD-1320, SRS-SEC-USB-1510
Test Execution Instructions
Save all captured screenshots and log files using the applicable test-step number for easy identification
Test environment
n/a
Definitions or Acronyms Used
n/a
Instruments & Tools
1. Minicom or PuTTY terminal software 2. High-speed USB-to-UART cable capable of supporting a baud rate of 1.5 Mbps 3. USB flash drive 4. A PC running a Linux operating system is recommended.
Fail/Pass Criteria
The test shall be considered passed if the actual results meet the expected results specified in the ‘Test Script’ sheet
Test Report Reference:
TBD
Comments:
The test executor shall have understanding of the device software upgrade using USB flash drive The test executor shall understand the use of minicom or equivalent serial terminal. The test executor shall be able to execute and read the Linux commands.

## Test Script
GTD-001228: Software Integration Test Request -06
Test Case No. | Test Descriptions | Expected Result | Actual Result
The following steps verify the displayed software and hardware version information.
10 | Power on the device and complete the initial configuration until the Welcome page is displayed. | The device powers on successfully and displays the Welcome page. | Did the device power on successfully and display the Welcome page? _____
20 | Navigate to the Device Information page. | The Device Information page opens successfully. | Did the Device Information page open successfully? _____
30 | Verify that the Device Information page displays the following version information: 1. OS Build Version 2. Application Version 3. Hardware Version 4. FPGA Version | The device displays the OS Build, Application, Hardware and FPGA version information without any missing values. | Did the device display all required version information? _____ Record the displayed versions: OS Build Version: _____ Application Version: _____ Hardware Version: _____ FPGA Version: _____
40 | Verify that the displayed version information matches the version information specified on the About sheet. | All displayed version information matches the version information specified on the About sheet. | Did all displayed version information match the About sheet? _____
Test ID: STS-ProVuVista-Integration-06-000010 Software Test: SRS-SEC-SUP-1030-05, SRS-SEC-SUP-1030-01, SRS-SEC-SUP-1870 The following steps verify that the software upgrade can handle the driver and kernel-level updates.
50 | Obtain the test software bundles supplied with this test method. | The required test software bundles are available for test execution. | Setup step completed. No verification response is required.
60 | Connect the USB-to-UART serial cable and configure the serial terminal baud rate to 1.5 Mbps. Log in with the username 'root' and a blank password. Run the following command to set the kernel log level: dmesg -n 2 | The serial connection is established, root login succeeds and the command is accepted without an error. | Did the serial connection establish successfully and did the root login succeed? _____ Did the command execute without errors? _____
70 | Power on the device. | The device powers on and completes the boot process successfully. | Did the device power on and complete the boot process successfully? _____
80 | Copy the upgrade file to the USB drive from software bundles provided with the test method under folder name 'UpgradeBundle-NoEEPROMDriver'. | The specified upgrade file is copied to the USB drive successfully. | Did the specified upgrade file copy to the USB drive successfully? _____
90 | Navigate to the Settings screen and verify that the Service menu remains unavailable until a user logs in. | The Service menu is unavailable before user login. | Did the Service menu remain unavailable before user login? _____
100 | Log in as a Standard user account (create if no account created) and log in as a standard user. | The Service menu remains unavailable to the Standard user. | Did the Service menu remain unavailable to the Standard user? _____
110 | Log in using the administrator credentials and create a service user account. Log out of the administrator account, then log in using the service user account. Connect the USB drive and initiate the software upgrade. Confirm the prompts until the installation confirmation prompt appears. Verify that the installation confirmation appears on the screen showing the version of the upgrade file. Record the software version and confirm the installation. | The Service user account is created successfully and the user can log in. The installation confirmation prompt displays the upgrade-file software version. | Did the Service user account get created successfully and did login succeed? _____ Did the installation confirmation prompt display the upgrade-file software version? _____ Record the displayed upgrade-file software version: _____
120 | Go to Event Logs and check for event ID 170 | The software shall log the Event with ID 170 and message "Software upgrade initiated" | Did software log the event? ________
120 | Confirm the installation prompt and verify that a progress bar showing installation status in percentage displayed. When the upgrade complete prompt appears, select 'OK' to restart the device. Note: Check "references" sheet for information over prompts | The installation progress bar displays the installation percentage. The device displays a “Success” message, records event ID 171 in the event log, and restarts after “OK” is selected. | Did the installation progress bar display the installation percentage? _____ Did the software upgrade complete successfully without errors? _____ Did the device display the “Success” message? _____ Did the event log record event ID 171? _____ Did the device restart after “OK” was selected? _____ Record the software version displayed during installation: _____
130 | Navigate to the Device Information page and verify that the displayed software version matches the version recorded during installation. | The software version displayed on the Device Information page matches the version recorded during installation. | Did the displayed software version match the version recorded during installation? _____ Record the displayed software version: _____
140 | Connect a bronchoscope containing an EEPROM and power cycle the device. Run the following command: i2cdetect -y 5 Verify that the output shows device address 0x18. Note: This step verifies that the bronchoscope EEPROM responds on I2C bus 5 at address 0x18. Because the driver is not installed, the EEPROM is not bound to the driver. | The command output displays the bronchoscope EEPROM address 0x18 on I2C bus 5. | Did the command output display EEPROM device address 0x18 on I2C bus 5? _____ Record the command output: _____
150 | Upgrade the device using the software bundle located in the folder named 'UpgradeBundle-ReleaseUSB'. | The device accepts the specified release bundle and starts the software upgrade. | Did the device accept the release bundle and start the upgrade? _____
160 | Complete the installation and record the software version displayed during the upgrade. Navigate to the Device Information page and compare the displayed version with the recorded installation version. | The installation completes successfully, and the version on the Device Information page matches the version recorded during installation. | Did the installation complete successfully? _____ Did the displayed version match the recorded installation version? _____ Record the installation version: _____ Record the Device Information version: _____
170 | Reconnect the bronchoscope containing EEPROM, power cycle the device and verify that the EEPROM device entry appears under the 1-Wire control path. Exact path: /sys/bus/w1/devices/2d-xxxxxx Note: The serial number appeared under devices can change, as it depends on the EEPROM unique ID | The EEPROM device is detected under the specified 1-Wire control path. | Did the EEPROM device appear under the specified 1-Wire control path? _____ Record the path and device entry: _____
Test ID: STS-ProVuVista-Integration-06-000020 Software Test: SRS-SEC-SUP-1610, SRS-SEC-SUP-1970, SRS-SEC-SUP-1590, SRS-SEC-SUP-1030-02 The following steps verify that the software upgrade can detect bundle tampering/modification.
180 | Copy the software upgrade file provided in the folder UpgradeBundle-InvalidSignature and attempt the software upgrade. Note: this step checks the software upgrade can detect if upgrade package modified during transfer | The device shall reject the upgrade bundle due to invalid signature. The device shall display an error prompt with message "The software upgrade package is corrupt. Please use correct software." The failure event shall be recorded in event log with event ID 173. | Did the device reject the invalid-signature bundle? _____ Did the device display the specified error message? _____ Did the event log record upgrade failure event ID 173? _____ Record the displayed message and event ID: _____
190 | Copy the software upgrade file provided in the folder UpgradeBundle-InvalidSizeHigh and attempt the software upgrade. Note: this step checks the software upgrade can detect if upgrade package modified during transfer | The device shall reject the upgrade bundle due to invalid size. The device shall display an error prompt with message "The software upgrade package is corrupt. Please use correct software." The failure event shall be recorded in event log with event ID 173. | Did the device reject the invalid-size bundle? _____ Did the device display the specified error message? _____ Did the event log record upgrade failure event ID 173? _____ Record the displayed message and event ID: _____
200 | Copy the software upgrade file provided in the folder UpgradeBundle-InvalidSizeLow and attempt the software upgrade. Note: this step checks the software upgrade can detect if upgrade package modified during transfer | The device shall reject the upgrade bundle due to invalid size. The device shall display an error prompt with message "The software upgrade package is corrupt. Please use correct software." The failure event shall be recorded in event log with event ID 173. | Did the device reject the invalid-size bundle? _____ Did the device display the specified error message? _____ Did the event log record upgrade failure event ID 173? _____ Record the displayed message and event ID: _____
210 | Copy the software upgrade file provided in the folder UpgradeBundle-IncorrectName and attempt the software upgrade. Note: this step checks the software upgrade can detect if upgrade package modified during transfer | The device shall reject the upgrade bundle due to invalid name. The device shall display an error prompt with message "The software upgrade package is corrupt. Please use correct software." The failure event shall be recorded in event log with event ID 173. | Did the device reject the incorrect-name bundle? _____ Did the device display the specified error message? _____ Did the event log record upgrade failure event ID 173? _____ Record the displayed message and event ID: _____
220 | Verify that device displays the required file name (update-bundle-provu.zip) to the user with error message. | The device displays the required file name, “update-bundle-provu.zip”, in the error message. | Did the error message display the required file name, “update-bundle-provu.zip”? _____ Record the displayed message: _____
Test ID: STS-ProVuVista-Integration-06-000030 Software Test: SRS-SEC-SUP-1030-03 The following steps verify that the software upgrade can detect the USB removal while upgrade package is copying into device.
230 | Copy the software upgrade file provided in the folder UpgradeBundle-ReleaseUSB and attempt the software upgrade. While the software displays the "download-prompt" then remove the USB drive | The device shall detect the USB media removal. The device shall show the error message "Error updating the software. Please try again or view Error Log." The device shall record the error event in log with ID 172 | Did the device detect removal of the USB drive? _____ Did the device display expected error message? _____ Record the displayed message: _____
240 | Power-cycle the device and verify that it completes the boot process. | The device boots successfully using the previously installed software. | Did the device boot successfully using the previously installed software? _____
250 | Initiate the software upgrade again. After the software has been downloaded and while the device displays the 'Installing software' prompt, remove the USB drive. Record the installation version displayed during upgrade. | Removal of the USB drive during installation does not interrupt the installation because the package has already been copied to the device. The upgrade completes successfully. | Did the installation continue after the USB drive was removed? _____ Did the software upgrade complete successfully? _____ Record the installation version: _____
260 | When the upgrade complete prompt appears, select 'OK' to restart the device. | The device restarts successfully and boots into the installed software version. | Did the device restart successfully and boot into the installed software version? _____ Record the installed version: _____
Test ID: STS-ProVuVista-Integration-06-000040 Software Test: SRS-SEC-SUP-1030-03 The following steps verify that the software upgrade can handle the power interruption.
270 | Remove the battery, power the device using the external power adapter and run the following command in the serial terminal: rauc status Record the currently active partition. Note: refer the sheet "references" for partition table | The command completes successfully and identifies the active partition. | Did the 'rauc status' command complete successfully? _____ Record the active partition: _____ Attach or record the command output: _____
280 | Using the same upgrade file provided in the folder name UpgradeBundle-ReleaseUSB, initiate the software upgrade. While the device displays the 'Installing software' prompt, interrupt power and then restore power to the device. | The device boots successfully after the interrupted installation. | Did the device boot successfully after the interrupted installation? _____
290 | Run the following command in the serial terminal and compare the partition status with the result recorded before the interrupted installation: rauc status | The previously active partition remains active, and the inactive partition targeted by the interrupted installation is marked as 'bad'. | Did the previously active partition remain active? _____ Did the target inactive partition become marked as 'bad'? _____ Record the partition status: _____
300 | Using the same upgrade file, initiate and complete another the software upgrade without interrupting power. | The software upgrade completes successfully and the device boots successfully. | Did the software upgrade complete successfully? _____ Did the device boot successfully? _____
310 | Run the following command in the serial terminal and verify the status of the partition previously marked as 'bad': rauc status | The partition previously marked as 'bad' is programmed successfully and becomes the active boot partition. | Did the partition previously marked as 'bad' programme successfully? _____ Did the partition become the active boot partition? _____ Record the partition status: _____
Test ID: STS-ProVuVista-Integration-06-000040 Software Test: SRS-SEC-SUP-1810 The following steps verify that the software upgrade can detect failure with a partition and switch to other partition after three attempts.
320 | During boot, interrupt power while the progress bar is displayed. Restore power and repeat the boot interruption three times. After the third interruption, restore power, allow the device to boot and run: rauc status | After three consecutive boot interruptions, the device switches to and boots from the secondary partition. | Did the device switch to the secondary partition after three boot interruptions? _____ Did the device boot successfully from the secondary partition? _____ Record the active partition and RAUC status: _____
Test ID: STS-ProVuVista-Integration-06-000050 Software Test: SRS-SEC-SUP-1030-04 The following steps verify that the software upgrade can be done if device battery is less than 20%.
330 | 1. Connect the battery and discharge it until the indicated battery level is below 20%. 2. Disconnect the power adapter from device Note: Do not discharge the battery below 5%, as after it the device enters ship mode | The device indicates a battery level below 20%. | Did the indicated battery level fall below 20%? _____ Record the battery level: _____ %
340 | With the battery level below 20%, initiate an upgrade using the release bundle provided in the folder name UpgradeBundle-ReleaseUSB. | The device permits the software upgrade at a battery level below 20%. | Did the device permit the software upgrade at a battery level below 20%? _____
350 | 1. Disconnect the battery I2C lines and power cycle the device. 2. Attempt the software upgrade using the release bundle provided in the folder “UpgradeBundle-ReleaseUSB”. | The device permits the software upgrade when the battery is faulty. | Did the device permit the software upgrade when the battery was faulty? _____
360 | 1. Connect a battery with a reported battery health below 70%. 2. Attempt the software upgrade using the release bundle provided in the folder “UpgradeBundle-ReleaseUSB”. | The device permits the software upgrade when the battery is degraded. | Did the device permit the software upgrade when the battery was degraded? _____
370 | 1. Connect the power adapter. 2. Remove the battery from the device. 3. Attempt the software upgrade using the release bundle provided in the folder “UpgradeBundle-ReleaseUSB”. | The device permits the software upgrade when the battery is not connected and the power adapter is connected. | Did the device permit the software upgrade when the battery was not connected and the power adapter was connected? _____
Test ID: STS-ProVuVista-Integration-06-000060 Software Test: SRS-SEC-SUP-1030-02, SRS-SEC-SUP-2190 The following steps verify that the software upgrade can detect the tampering with the bundle contents.
380 | Copy the software upgrade file provided in the folder UpgradeBundle-InvalidManifest and attempt the software upgrade. | The device shall reject the upgrade bundle. The device shall display an error prompt with message "The software upgrade package is corrupt. Please use correct software." The failure event shall be recorded in event log with event ID 173. | Did the device reject the invalid-manifest bundle? _____ Did the device display the specified error message? _____ Did the event log record upgrade failure event ID 173? _____ Record the displayed message and event ID: _____
Test ID: STS-ProVuVista-Integration-06-000070 Software Test: SRS-SEC-SUP-2190 The following steps verify that the software upgrade package contains a list of compatible hardware, software versions and upgrade bundle file size.
390 | Extract the software bundle and open the manifest file using a text editor. | The bundle is extracted successfully and the manifest file can be opened. | Did the bundle extract successfully? _____ Did the manifest file open successfully? _____
400 | Verify that the manifest file contains the 1. Hardware PCB IDs (provu_pcb_ids) 2. Software Version (bundle_version) 3. Upgrade File Size (bundle_size) | The manifest contains all the specified parameters | Did the manifest contain all required fields? _____ Record the PCB ID: _____ Record the software version: _____ Record the upgrade file size: _____
Test ID: STS-ProVuVista-Integration-06-000080 Software Test: SRS-EXT-ITF-1490, SRS-SEC-SUP-1990 The following steps verify that the software upgrade notifies the user for incompatible hardware and software versions but allow user to upgrade the device.
410 | Copy the software upgrade file provided in the folder UpgradeBundle-InvalidPCB-ID and attempt the software upgrade. | The device shall not reject the upgrade and upgrade shall complete successfully. The device shall display an error prompt with message "The software upgrade package failed. Software package not compatible with hardware version." The failure event shall be recorded in event log with event ID 174. | Did the software upgrade complete successfully? _____ Did the event log record the upgrade failure with event ID 174 ? _____ Record the displayed message and event identifier: _____
420 | Copy the software upgrade file provided in the folder UpgradeBundle-OlderSoftwarePackage and attempt the software upgrade. | The device shall not reject the upgrade and upgrade shall complete successfully. The device shall display an error prompt with message "Software update package is older. Please use correct software." The failure event shall be recorded in event log with event ID 176. | Did the screen recording save successfully? _____ Record the file name: _____ Did the Standard user account get created successfully? _____ Record an event ID: _____ Record an error ID: _____
Test ID: STS-ProVuVista-Integration-06-000090 Software Test: SRS-SEC-SUP-1800 The following steps verify that the software upgrade programmes only the inactive root filesystem partition and does not modify the data partition.
430 | 1. Create and save a screen recording, record its file name 2. Create a standard user account. 3. Record the Event log IDs (any one of them) 4. Record the Error log IDs (any one of them) | The screen recording is saved, a Standard user account is created, and the selected event and error log IDs are recorded. | Did the screen recording save successfully? _____ Record the file name: _____ Did the Standard user account get created successfully? _____ Record an event ID: _____ Record an error ID: _____
440 | 1. Log in as Admin user 2. Change the Image control brightness to maximum for bronchoscope. 3. Save the configuration and logout the Admin | The bronchoscope image-control brightness is set to the maximum value and saved successfully for the Administrator account. | Did the bronchoscope image-control brightness change to the maximum value? _____ Did the Administrator account retain the saved brightness setting after logout and login? _____
440 | 1. Log in as Standard User 2. set the color adjustment to maximum value for VL camera 3. Save the configuration and logout | The VL camera colour adjustment is set to the maximum value and saved successfully for the Standard user account. | Did the VL camera colour adjustment change to the maximum value? _____ Did the Standard user account retain the saved colour setting after logout and login? _____
450 | Create a service user account if not created. | The Service user account is created successfully. | Did the Service user account get created successfully? _____
460 | Copy the release bundle to a USB drive and complete the software upgrade. After the upgrade, verify that: (a) the previously saved screen recording remains available; (b) all existing event and error logs remain available; and (c) the standard user can log in using the account created before the upgrade and color adjustment change for standard user remains same (d) Login as Admin and check image adjustment done by admin account are saved (e) Login as service user and check service account can be accessed. | The upgrade completes successfully. The saved recording, event logs, error logs, user accounts, and image-control settings remain available and unchanged after the upgrade. | Did the upgrade complete successfully? _____ Did the saved recording remain available? _____ Did the existing event and error logs remain available? _____ Did the Standard user log in successfully? _____ Did the Administrator user log in successfully? _____ Did the Service user log in successfully? _____ Did the Standard and Administrator user image-control settings remain unchanged? _____ Record evidence or file references: _____
Test ID: STS-ProVuVista-Integration-06-000100 Software Test: SRS-SEC-SUP-1620 The following steps verify that power button is disabled during the software upgrade.
470 | 1. Copy the software upgrade file provided in the folder UpgradeBundle-ReleaseUSB and attempt the software upgrade. 2. While the software upgrade in progress, press the power button and check device does not enter sleep mode or power off | The device remains powered on and does not enter sleep mode while the software upgrade is in progress. | Did the device remain powered on and avoid entering sleep mode while the software upgrade was in progress? _____
480 | During the upgrade, press and hold the power button for 10 seconds. Note: Checking that if the software upgrade becomes unresponsive then using power button device can be restarted forcibly | The device shall restart successfully | Did the 10-second power-button hold restart the device? _____ Did the device subsequently boot successfully? _____ Record the observed power action: _____
Test ID: STS-ProVuVista-Integration-06-000110 Software Test: SRS-SEC-SUP-2190 The following steps verify that the software upgrade can handle if manifest file modified and checksum updated as per new contents.
490 | Copy the software upgrade file provided in the folder UpgradeBundle-InvalidManifest-CorrectSha256 and attempt the software upgrade. Note: this bundle updates the manifest contents and place the correct SHA-256 of the modified contents. | The device shall reject the upgrade bundle. The device shall display an error prompt with message "The software upgrade package is corrupt. Please use correct software." The failure event shall be recorded in event log with event ID 173. | Did the device reject the modified-manifest bundle with the updated SHA-256 checksum? _____ Did the device display the specified error message? _____ Did the event log record upgrade failure event ID 173? _____ Record the displayed message and event ID: _____
Test ID: STS-ProVuVista-Integration-06-000120 Software Test: SRS-SEC-HARD-1320 The following steps verify that the root file system is mounted as read-only
500 | Copy the upgrade file provided in the folder name UpgradeBundle-ReleaseUSB to USB drive and program the device. | The software upgrade completes successfully and the device boots successfully. | Did the software upgrade complete successfully? _____ Did the device boot successfully? _____ Record the software version displayed under information screen ________
510 | Using the serial terminal, run the following commands echo "Hello World" > /tmp/testfile.txt ls //list files Note: // defines the comments not part of command | The file is created successfully in /tmp. | Did the file get created successfully in /tmp? _____ Record the command output: _____
520 | Now run the following commands and record output cp -f /tmp/testfile.txt /usr/bin cp -f /tmp/testfile.txt /usr/ cp -f /tmp/testfile.txt /etc/ cp -f /tmp/testfile.txt /bin/ cp -f /tmp/testfile.txt /usr/lib/ cp -f /tmp/testfile.txt /lib/ | Each copy command fails because the target root file system locations are read-only. | Did each copy command fail? _____ Record the command outputs: _____
Test ID: STS-ProVuVista-Integration-06-000130 Software Test: SRS-SEC-SUP-1980 The following steps verify the eMMC partitions
530 | From serial terminal, run command lsblk | The command output shall display the current eMMC (mmc0) partitions | Did the command output show the eMMC (mmc0) partition blocks? _____ Record the command output: _____
540 | Verify that below partitions are available Part0: Label: u-boot Part1: Label: rootfsA Part2: Label: rootfsB Part3: Label: userdata | The specified partitions shall be available | Did the command output list all specified partitions and labels? _____ Record the command output: _____
Test ID: STS-ProVuVista-Integration-06-000140 Software Test: SRS-SEC-USB-1510 The following steps verify the software allows only camera devices over micro-USB ports and storage devices on USB-A port
550 | Connect USB mouse Bluetooth dongle to the USB-A port | The device does not detect the mouse, and the mouse remains unusable. | Did the mouse remain undetected and unusable when connected to the USB-A port? _____
560 | Connect USB mouse Bluetooth dongle with the Micro-USB port, with the help of a USB-A to Micro-USB port converter | The device does not detect the mouse, and the mouse remains unusable. | Did the mouse remain undetected and unusable when connected to the Micro-USB port? _____
570 | Connect a USB-A camera to the USB-A port. | The device does not display the camera feed when the USB-A camera is connected to the USB-A port. | Did the device prevent the camera feed from being displayed? _____
580 | Connect a USB flash drive to the Micro-USB camera port. | The device does not display the USB storage-device icon when the USB flash drive is connected to the Micro-USB camera port. | Did the device prevent the USB storage-device icon from being displayed in the top-right corner? _____

## references
download-prompt
installation-prompt
installation-confirmation
upgrade-complete