## CoverPage
155 - ProVu Accessories
ProVu Vista Display – Software Integration Test Request - 09
Project: |  | 155 - ProVu Accessories
Document ID/Version: |  | GTD-001263/A
DOCUMENT HISTORY
Version | Document Changes |  | Author | Date
A | Initial release |  | Ayush Pareek | 2026-09-22
Associated Documents
Document No |  | Document Title
SRS-000053 |  | ProVu Vista Display –SOFTWARE REQUIREMENT SPECIFICATION
GTD-000993 |  | PROVU2 LARGE DISPLAY Software Test Plan
Attachment
Name |  | Description And Summary
n/a |  | n/a

## About
GTD-001263: Software Integration Test Request -09
Test Request Information
Author | Ayush Pareek
Revision | A
Reference | GTD-000993 | Date | 2026-09-22
Test objective
This test method verifies the - Warning and Errors - Internal Storage Full
Test Coverage
SRS-FUN-ENW-990, SRS-FUN-STG-1020
Test Execution Instructions
n/a
Test environment
Linux PC with minicom installed to talk over serial terminal or Else a Window Host PC with putty
Definitions or Acronyms Used
n/a
Instruments & Tools
1. Linux or Windows PC 2. USB to UART cable (high speed 1.5Mbps)
Fail/Pass Criteria
The test shall be considered passed when the actual result for each applicable test step meets the corresponding expected result in the "Test Script" sheet.
Test Report Reference:
TBD
Comments:
1. If a test step require to configure the serial terminal then configure the baud rate to 1.5Mbps 2. Once device boot up then login using "username" as root. Leave password blank 3. Set kernel log level to 2 by running command "dmesg -n 2"

## Test Script
GTD-001263: Software Integration Test Request -09
Test Case No. | Test Descriptions | Expected Result | Actual Result
The following steps verify the software version information.
10 | Power on the device. Configure the language, date and time, and user account. | The device configuration shall be completed successfully. | Did the device configuration complete successfully? _____ Record the configured username: _____ Record the configured password: _____
20 | Open Settings and navigate to the Information screen. | The Information screen shall open successfully. | Did the Information screen open successfully? _____
30 | Verify that the Information screen displays the following version information: 1. OS build version 2. Application version 3. Hardware version 4. FPGA version | The device shall display all specified version information. | Did the device display all specified version information successfully? _____ Record the OS build version: _____ Record the application version: _____ Record the hardware version: _____ Record the FPGA version: _____
40 | Compare the displayed version information with the target Software Release Form. | All displayed versions shall match the versions specified in the target Software Release Form. | Did all displayed versions match the target release versions? _____ Record the Software Release Form reference: _____ Record any mismatch: _____
Test ID: STS-ProVuVista-Integration-09-000010 Software item under test: SRS-FUN-ENW-990 This section verifies the event related to USB Port
50 | 1. Connect a USB flash drive to the device 2. Go to user event logs under service menu | The software shall log the event with ID 00001. The event message shall be "USB storage device connected" | Did the software log the specified entry with ID 00001?____ Did the event message match "USB storage device connected"? ________
60 | Select the "Export" button to transfer event logs to USB drive. | The software shall display a "Prompt pop-up" with message "File transfer complete" | Did the software display the specified pop-up message "File transfer complete"?___________
70 | Go to user event logs under service menu | The software shall log the data-transfer event with ID 00002. The event message shall be "Data transferred to USB storage device successfully" | Did the software log the specified event with ID 00002? _______ Did the event message match "Data transferred to USB storage device successfully"? __________
80 | Disconnect the USB flash drive from the device | The software shall log the event with ID 00004. The event message shall be "USB storage device disconnected" | Did the software log the specified entry with ID 00004?_______ Did the event message match "USB storage device disconnected"? ____________
90 | Connect a USB camera to the device through USB-A port. | The software shall display a "Warning pop-up" with message "Unsupported USB device." | Did the software display the specified pop-up message "Unsupported USB device."? ___________
100 | Go to error logs under service menu | The software shall log the error with ID 00005. The Error message shall be "Invalid USB class device attached" | Did the software log the specified entry with ID 00005? ______________ Did the observed result meet the following expected result: The Error message shall be "Invalid USB class device attached"? __________
Test ID: STS-ProVuVista-Integration-09-000020 Software item under test: SRS-FUN-ENW-990 This section verifies the event related to Camera Ports
110 | 1. Connect a bronchoscope camera to the device 2. Go to user event logs under service menu | The software shall log the event with ID 00020. The event message shall be "Camera connected to port#1" | Did the software log the specified entry with ID 00020? ____________ Did the event message match "Camera connected to port#1"? _______
120 | 1. Disconnect the bronchoscope camera from the device 2. Go to user event logs under service menu | The software shall log the event with ID 00021. The event message shall be "Camera disconnected to port#1" | Did the software log the specified entry with ID 00021? ____________ Did the event message match "Camera disconnected to port#1"? _____________
130 | 1. Connect a USB camera to Port B of the device 2. Go to user event logs under service menu | The software shall log the event with ID 00022. The event message shall be "Camera connected to port#2" | Did the software log the specified entry with ID 00022? _____________ Did the event message match "Camera connected to port#2"? _____________
140 | 1. Disconnect the USB camera from Port B of the device 2. Go to user event logs under service menu | The software shall log the event with ID 00023. The event message shall be "Camera disconnected to port#2" | Did the software log the specified entry with ID 00023? _________________ Did the event message match "Camera disconnected to port#2"? _________________
150 | 1. Connect a USB camera to Port C of the device 2. Go to user event logs under service menu | The software shall log the event with ID 00024. The event message shall be "Camera connected to port#3" | Did the software log the specified entry with ID 00024? ______________ Did the event message match "Camera connected to port#3"? ____________
160 | 1. Disconnect the USB camera from Port C of the device 2. Go to user event logs under service menu | The software shall log the event with ID 00025. The event message shall be "Camera disconnected to port#3" | Did the software log the specified entry with ID 00025? ___________ Did the event message match "Camera disconnected to port#3"? __________________
170 | Connect a faulty bronchoscope camera the device | The software shall display a "Warning pop-up" with message Internal hardware problem on camera port 1. Service required by local representative. | Did the software display the specified pop-up message ? __________ Did the observed result meet the following expected result? __________
180 | Connect a faulty USB camera to Port B of the device | The software shall display a "Warning pop-up" with message Internal hardware problem on camera port 2. Service required by local representative. | Did the software display the specified pop-up message ? __________ Did the observed result meet the following expected result? _________
190 | Connect a faulty USB camera to Port C of the device | The software shall display a "Warning pop-up" with message Internal hardware problem on camera port 3. Service required by local representative. | Did the software display the specified pop-up message ? ________ Did the observed result meet the following expected result? ________
200 | Go to error logs under service menu | The software shall log the error with ID 00026, 00027, 00028 and message specified in "Error Event Table column "Error/ Event Description/ Log text" | Did the software log the specified entry with ID 00026, 00027, 00028 and message specified in "Error Event Table column "Error/ Event Description/ Log text"? _________
Test ID: STS-ProVuVista-Integration-09-000030 Software item under test: SRS-FUN-ENW-990 This section verifies the events related to Standard and Admin User accounts
210 | 1. Login as Admin and go to user account creation page 2. Enter different fields on "Password" and "Re-type password" and confirm | The software shall display a "Message pop-up" with message "Passwords do not match." | Did the software display the specified pop-up message "Passwords do not match."? __________
220 | Leave "Username" field empty and confirm the account creation | The software shall display a "Message pop-up" with message "Please complete all fields." | Did the software display the specified pop-up message "Please complete all fields."? ___________
230 | Enter "Username" as AP@123 | The software shall display a "Message pop-up" with message "Username must not contain invalid characters." | Did the software display the specified pop-up message "Username must not contain invalid characters."? ______
240 | Enter Password as Ap@12 | The software shall display a "Message pop-up" with message "Password must be at least 8 characters long." | Did the software display the specified pop-up message "Password must be at least 8 characters long."? ________
250 | Enter the details below and create account First name: Flexi Last name: Care Username: Flexicare Password: Flexi@12345 Re-type password: Flexi@12345 Do not select account type | The software shall display a "Message pop-up" with message "Please select an account type (Service or Standard)." | Did the software display the specified pop-up message "Please select an account type (Service or Standard)."? __________
260 | Select Account Type as Standard | The user account shall be created successfully. | Did the user account be created successfully? _________
270 | Log in as Admin and go to account creation page. Enter below details First name: Flexi Last name: Care Username: Flexicare Password: Flexi@12345 Re-type password: Flexi@12345 Account Type: Service | The software shall display a "Message pop-up" with message "This username is already in use. Please choose a unique username." | Did the software display the specified pop-up message "This username is already in use. Please choose a unique username."? _____________
280 | 1. Log in to the device using username "Flexicare" and Password "Flexi@12345" 2. Go to Account creation and update password for user "Flexicare" as "Flexi@123456" | The software shall display a "Message pop-up" with message "Password updated successfully." The software shall display a "Message pop-up" with message "Account details updated." | Did the software display the specified pop-up message "Password updated successfully."? ___________ Did the software display the specified pop-up message "Account details updated."? ________
290 | Log out and try to log in using the old password "Flexi@12345" | Login using the old password shall fail. | Did login using the old password fail? __________
300 | Log in as Flexicare user | The software shall display a "Message pop-up" with message "User logged in successfully." | Did the software display the specified pop-up message "User logged in successfully."? __________
310 | Go to user event logs under service menu | he software shall log the event with ID 00050. The event message shall be "User <x> logged in ( x--> User Name)" | Did the observed result meet the following expected result: he software shall log the event with ID 00050? _______ Did the event message match "User <x> logged in ( x--> User Name)"? ____________
320 | Log out from Flexicare user | The software shall display a "Message pop-up" with message "User logged out." | Did the software display the specified pop-up message "User logged out."? ____________
330 | Go to user event logs under service menu | The software shall log the event with ID 00051. The event message shall be "User <x> logged out ( x--> User Name)" | Did the software log the specified entry with ID 00051? Did the event message match "User <x> logged out ( x--> User Name)"? _________
340 | Log in to the device as Admin and change the admin account password to "Admin@9462" | The software shall display a "Message pop-up" with message "Admin logged in" The software shall display a "Message pop-up" with message "Admin password updated" | Did the software display the specified pop-up message "Admin logged in"? ________________ Did the software display the specified pop-up message "Admin password updated"? __________
350 | Log out from Admin account | The software shall display a "Message pop-up" with message "Admin logged out" | Did the software display the specified pop-up message "Admin logged out"? ___________
360 | Log in as Admin and create a user account with username "Flexi0001" | The software shall display a "Message pop-up" with message "New user account created" | Did the software display the specified pop-up message "New user account created"? _________
370 | Reset the user "Flexi0001" password | The software shall display a "Message pop-up" with message "Password reset successfully." | Did the software display the specified pop-up message "Password reset successfully."? ____________
380 | Delete the user "Flexi0001" account | The software shall display a "Message pop-up" with message "User account deleted." | Did the software display the specified pop-up message "User account deleted."? __________
390 | 1. Go to the event logs under service menu 2. Verify the event logged from ID 50 to 75 with the message specified in "Error Event Table" column "Error/ Event Description/ Log text" | All specified event IDs shall be available with the corresponding messages listed in the "Error Event Table". | Did all specified event IDs appear with the corresponding messages listed in the "Error Event Table"? _________
Test ID: STS-ProVuVista-Integration-09-000040 Software item under test: SRS-FUN-ENW-990 This section verifies the events related to Service User account
400 | 1. Login as Admin and create a service user account with username "service01" 2. Log out from Admin account and log in as service user | The software shall display a "Message pop-up" with message "Service user logged in." | Did the software display the specified pop-up message "Service user logged in."?___________
410 | Log out from service user | The software shall display a "Message pop-up" with message "Service user logged out." | Did the software display the specified pop-up message "Service user logged out."? ___________
420 | Log in as service user and change password | The software shall display a "Message pop-up" with message "Service password updated." | Did the software display the specified pop-up message "Service password updated."? __________
430 | 1. Connect USB flash drive to the device 2. Go to error logs and export it | The software shall display a "Message pop-up" with message "error logs exported successfully." | Did the software display the specified pop-up message "error logs exported successfully." ? ____________
440 | Go to event logs and export it | The software shall display a "Message pop-up" with message "event logs exported successfully." | Did the software display the specified pop-up message "event logs exported successfully."? ______________
450 | Export Event log and immediately remove USB media | The software shall display a "Message pop-up" with message "Error exporting Error Logs. Please try again or view error log." | Did the software display the specified pop-up message "Error exporting Error Logs. Please try again or view error log."? ___________
460 | Export Event log and immediately remove USB media | The software shall display a "Message pop-up" with message "Error exporting Error Logs. Please try again or view error log." | Did the software display the specified pop-up message "Error exporting Error Logs. Please try again or view error log."? ___________
470 | 1. Go to the event logs under service menu 2. Verify the events 90, 91, 92, 93, 94, 95, 96 with the message specified in "Error Event Table" column "Error/ Event Description/ Log text" | All specified event IDs shall be available with the corresponding messages listed in the "Error Event Table". | Did all specified event IDs appear with the corresponding messages listed in the "Error Event Table"?_________
Test ID: STS-ProVuVista-Integration-09-000050 Software item under test: SRS-FUN-ENW-990 This section verifies the events related to Media Files
480 | 1. Record screen for 2 min 2. Take 5 screen image capture 3. Go to device gallery and delete the recording file | The software shall display a "Message pop-up" with message "Selected files deleted." | Did the software display the specified pop-up message "Selected files deleted."? _______
490 | Delete all the files including the screen image capture | The software shall display a "Message pop-up" with message "All media files deleted." | Did the software display the specified pop-up message "All media files deleted."? __________
500 | 1. Go to the event logs under service menu 2. Verify the events 111, 112 with the message specified in "Error Event Table" column "Error/ Event Description/ Log text" | All specified event IDs shall be available with the corresponding messages listed in the "Error Event Table". | Did all specified event IDs appear with the corresponding messages listed in the "Error Event Table"? _________
Test ID: STS-ProVuVista-Integration-09-000060 Software item under test: SRS-FUN-ENW-990 This section verifies the events related to Device Settings
510 | 1. Login as Admin and change the device date 2. Go to event log | The software shall log the event with ID 130 and message "RTC changed" The software shall log the event with ID 134 and message "Date & Time Updated " | Did the software log the specified entry with ID 130 and message "RTC changed"? _______ Did the software log the specified entry with ID 134 and message "Date & Time Updated "? _________
520 | Go to Admin menu and perform the following actions 1. Change Language 2. Change Maximum Video recording length 3. Change Power saving time | The software shall log the event with ID 131 and message "Language setting changed" The software shall log the event with ID 132 and message "Maximum Video recording length updated" The software shall log the event with ID 133 and message "Power saving time updated" | Did the software log the specified entry with ID 131 and message "Language setting changed"? ____________ Did the software log the specified entry with ID 132 and message "Maximum Video recording length updated"? _____________ Did the software log the specified entry with ID 133 and message "Power saving time updated"? _________
Test ID: STS-ProVuVista-Integration-09-000070 Software item under test: SRS-FUN-ENW-990 This section verifies the events related to Battery
530 | Connect a battery to the device with I2C lines disconnected | The software shall display a "Warning pop-up" with message Battery fault detected. Service required by local representative. | Did the software display the specified pop-up message ? ____________ Did the observed result meet the following expected result? ____________
540 | Connect a battery to the device with power line disconnected | The software shall display a "Warning pop-up" with message Battery charging fault. Check the power source. | Did the software display the specified pop-up message ? ____________ Did the observed result meet the following expected result? ____________
550 | Connect a battery to the device with NTC wire disconnected | The software shall display a "Warning pop-up" with message Internal hardware problem detected. Service required by local representative. | Did the software display the specified pop-up message ? _____________ Did the observed result meet the following expected result? ____________
560 | Go to error log and check for error ID 150 | The software shall log the error with ID 150 and message "Battery fault" | Did the software log the specified entry with ID 150 and message "Battery fault"? _____________
570 | Go to error log and check for error ID 155 | The software shall log the error with ID 155 and message "Battery charging fault" | Did the software log the specified entry with ID 155 and message "Battery charging fault"? __________
580 | Go to error log and check for error ID 159 | The software shall log the error with ID 159 and message "Battery secondary thermistor out of range." | Did the software log the specified entry with ID 159 and message "Battery secondary thermistor out of range."? ____________
Test ID: STS-ProVuVista-Integration-09-000080 Software item under test: SRS-FUN-ENW-990 This section verifies the events related to Device Status
590 | Press and hold the power button to reset the device | The device shall reset when the user presses and holds the power button. | Did the device reset when the user presses and holds the power button?___________
600 | Press the power button and enter sleep mode | The system shall enter sleep mode. | Did the system enter sleep mode? _________
610 | Press the power button to wake the device from sleep mode | The system shall exit sleep mode. | Did the system exit sleep mode? _______
620 | 1. Press the power button and enter sleep mode 2. Wait for 20 minutes for device to enter into power off | The device shall automatically shut down after remaining in sleep mode for 20 minutes. | Did the device automatically shut down after remaining in sleep mode for 20 minutes? __________
630 | Log in as service user and perform factory reset | The device shall complete the factory reset successfully. | Did the device complete the factory reset successfully?__________
640 | Go to event logs and check for events 203, 204, 206, 207, 209, 256 | The software shall record all the event IDs with the message specified in "Error Event Table" in column "Error/ Event Description/ Log text" | Did the software record all the event IDs with the message specified in "Error Event Table" in column "Error/ Event Description/ Log text"? _______________
650 | 1. Connect the serial debug cable to the device and configure it to access the terminal 2. Run commands . /etc/profile pkill provu | The device shall restart successfully. | Did the device restart successfully? _____________
660 | Go to error logs under service menu | The software shall log the error with ID 205 and message "Watchdog reset occurred" The software shall log the error with ID 208 and message "Application has exited" | Did the software log the specified entry with ID 205 and message "Watchdog reset occurred"? __________ Did the software log the specified entry with ID 208 and message "Application has exited"? _________
Test ID: STS-ProVuVista-Integration-09-000090 Software item under test: SRS-FUN-ENW-990 This section verifies the events related to video recording
670 | Start screen recording and press gallery button | The software shall display a "Message pop-up" with message "To view the gallery, you must first stop the recording." | Did the software display the specified pop-up message "To view the gallery, you must first stop the recording."? __________
680 | Press the Settings button | The software shall display a "Message pop-up" with message "To edit device settings, you must first stop the recording." | Did the software display the specified pop-up message "To edit device settings, you must first stop the recording."?__________
690 | Connect a battery charged below 20% | The software shall display a "Message pop-up" with message "Battery charge critical. Unable to start recording. Please plug in charger." | Did the software display the specified pop-up message? ___________ Did the observed result meet the following expected result? ___________
700 | 1. Set recording maximum length to 30 min if not set 2. start screen recording and wait for 30 min | The software shall display a "Message pop-up" with message "Video was split into multiple files." The software shall display a "Message pop-up" with message "Recording ended & Video saved. Maximum 30 min video length reached" | Did the software display the specified pop-up messages? __________ Did the observed result meet the following expected result "Video was split into multiple files."?_______ Did the observed result meet the following expected result "Recording ended & Video saved. Maximum 30 min video length reached"? _____________
710 | 1. Set recording maximum length to 60 min if not set 2. start screen recording and wait for 60 min | The software shall display a "Message pop-up" with message "Video was split into multiple files." The software shall display a "Message pop-up" with message "Recording ended & Video saved. Maximum 60 min video length reached" | Did the software display the specified pop-up message ? ___________ Did the observed result meet the following expected result: "Video was split into multiple files."? ________ Did the observed result meet the following expected result: "Recording ended & Video saved. Maximum 60 min video length reached"?_____________
720 | Go to event logs and check for events 250, 251, 252, 254, 255, 256 | The software shall record all the event IDs with the message specified in "Error Event Table" in column "Error/ Event Description/ Log text" | Did the software record all the event IDs with the message specified in "Error Event Table" in column "Error/ Event Description/ Log text"? _______________
Test ID: STS-ProVuVista-Integration-09-0000100 Software item under test: SRS-FUN-ENW-990 This section verifies the max account limit reached event ID 76
730 | 1. Login as Admin and create 50 standard user accounts. 2. Create another user account after 50 accounts | The software shall display a "Message pop-up" with message "Maximum number of 50 users reached." | Did the software display the specified pop-up message ? ________________ Did the observed result meet the following expected result: "Maximum number of 50 users reached."? __________
740 | Go to event log under service menu | The software shall log the event with ID 76 and message "Maximum number of user accounts reached" | Did the software log the specified entry with ID 76 and message "Maximum number of user accounts reached"? ______________
Test ID: STS-ProVuVista-Integration-09-0000100 Software item under test: SRS-FUN-ENW-990, SRS-FUN-STG-1020 This section verifies the low disk space error
750 | 1. Go to service menu and perform factory reset 2. Start screen recording with any of two camera devices plugged in 3. Wait for a low disk space error occur | The software shall display a "Warning pop-up" with message "Low disk space. Less than 30mins recording time remaining. Please delete recordings from the gallery to provide more space." | Did the software display the specified pop-up message ? ________________ Did the error message have text as expected? ________
760 | Go to error log and check for error ID 191 | The software shall log the error with ID 191 and message "Low disk space" | Did the software log the specified entry with ID 191 and message "Low disk space"? _____________

## Error Event Table
Event ID | Event category | Error/ Event Description/ Log text | Event Logs | Error Logs | GUI warning/Error Type | User notifications message | GUI screen Identifier
GTD-000989
1 | USB port | USB storage device connected | Yes | NO | Info | N/A | N/A
2 |  | Data transferred to USB storage device successfully | Yes | NO | Prompt POP-UP + info | File transfer complete | 107
3 |  | Data failed to transfer to USB storage device | NO | Yes | Prompt POP-UP + info | Error exporting from Gallery. Please try again or view Error Log. | 120
4 |  | USB storage device disconnected | Yes | NO | Info | N/A | N/A
5 |  | Invalid USB class device attached | NO | Yes | Warning POP-UP | Unsupported USB device. | N/A
6 |  | Insufficient space on USB storage device | NO | Yes | Warning POP-UP | Insufficient space on USB storage device | N/A
7 |  | USB storage device disconnected while copying | NO | Yes | Warning POP-UP | Error Exporting from gallery. | N/A
USB device removed during transfer.
8 |  | USB Storage device has duplicates files/folders | No | Yes | Prompt POP-UP + info | USB storage device has duplicate file or folder with name <file or folder name> | N/A
20 | Camera ports | Camera connected to port#1 | Yes | NO | Info | N/A | N/A
21 |  | Camera disconnected from port#1 | Yes | NO | Info | N/A | N/A
22 |  | Camera connected to port#2 | Yes | NO | Info | N/A | N/A
23 |  | Camera disconnected from port#2 | Yes | NO | Info | N/A | N/A
24 |  | Camera connected to port#3 | Yes | NO | Info | N/A | N/A
25 |  | Camera disconnected from port#3 | Yes | NO | Info | N/A | N/A
26 |  | Fault with camera port#1 (Display hardware) | NO | Yes | Warning POP-UP | Internal hardware problem on camera port 1. | N/A
Service required by local representative.
27 |  | Fault with camera port#2 (Display hardware) | NO | Yes | Warning POP-UP | Internal hardware problem on camera port 2. | N/A
Service required by local representative.
28 |  | Fault with camera port#3 (Display hardware) | NO | Yes | Warning POP-UP | Internal hardware problem on camera port 3. | N/A
Service required by local representative.
29 |  | Fault with camera port#1 ( Accessories ) | NO | Yes | Warning POP-UP | Fault detected on camera Port 1. | N/A
Check camera and connections.
30 |  | Fault with camera port#2 ( Accessories ) | NO | Yes | Warning POP-UP | Fault detected on camera Port 2. | N/A
Check camera and connections.
31 |  | Fault with camera port#3 ( Accessories ) | NO | Yes | Warning POP-UP | Fault detected on camera Port 3. | N/A
Check camera and connections.
32 |  | Fault with Bronchoscope ISP | NO | Yes | Warning POP-UP | Internal hardware problem on camera port 1. | N/A
Service required by local representative.
33 |  | 3rd camera is attached ( Two Camera already connected) | Yes | NO | Message POP-UP (if two camera already connected) | Maximum of 2 views at a time. Toggle one off to continue. | 46
34 |  | Bronchoscope analog ID invalid | NO | Yes | Warning POP-UP | Fault detected on camera Port 1. | N/A
Check camera and connections.
35 |  | USB Host#1 supply rail invalid | NO | Yes | Warning POP-UP | Fault detected on camera Port 2. | N/A
Check camera and connections.
36 |  | USB Host#2 supply rail invalid | NO | Yes | Warning POP-UP | Fault detected on camera Port 3. | N/A
Check camera and connections.
50 | User accounts | User <x> logged in ( x--> User Name) | Yes | NO | Message POP-UP | User logged in successfully. | N/A
51 |  | User <x> logged out ( x--> User Name) | Yes | NO | Message POP-UP | User logged out. | N/A
52 |  | User <x> changed password ( x--> User Name) | Yes | NO | Message POP-UP | Password updated successfully. | N/A
53 |  | User <x> Failed password entry ( x--> User Name) | Yes | NO | Info | N/A | N/A
54 |  | User <x> Account details successfully updated ( x--> User Name) | Yes | NO | Message POP-UP | Account details updated. | N/A
55 |  | Passwords and Re-Type Passwords not match. | Yes | NO | Message POP-UP | Passwords do not match. | 80
56 |  | Text box fields left empty. | Yes | NO | Message POP-UP | Please complete all fields. | 38
57 |  | User name already exists. | Yes | NO | Message POP-UP | This username is already in use. Please choose a unique username. | N/A
58 |  | User entered invalid characters in username. | Yes | NO | Message POP-UP | Username must not contain invalid characters. | N/A
59 |  | Password entered less than 8 characters. | Yes | NO | Message POP-UP | Password must be at least 8 characters long. | N/A
60 |  | Invalid password used. | Yes | NO | Message POP-UP | Password matches a restricted list and cannot be used. | N/A
61 |  | Account type not selected. | Yes | NO | Message POP-UP | Please select an account type (Service or Standard). | N/A
70 | Admin account | Admin logged in | Yes | NO | Message POP-UP | Admin logged in | N/A
71 |  | Admin logged out | Yes | NO | Message POP-UP | Admin logged out | N/A
72 |  | Admin changed password | Yes | NO | Message POP-UP | Admin password updated | N/A
73 |  | User <x> account created ( x--> User Name) | Yes | NO | Message POP-UP | New user account created. | N/A
74 |  | User <x> account deleted ( x--> User Name) | Yes | NO | Message POP-UP | User account deleted. | N/A
75 |  | User <x> reset password ( x--> User Name) | Yes | NO | Message POP-UP | Password reset successfully. | N/A
76 |  | Maximum number of user accounts reached | Yes | NO | Message POP-UP | Maximum 50 number of users reached. | N/A
90 | Service user | Service user logged in | Yes | NO | Message POP-UP | Service user logged in. | N/A
91 |  | Service user logged out | Yes | NO | Message POP-UP | Service user logged out. | N/A
92 |  | Service user changed password | Yes | NO | Message POP-UP | Service password updated. | N/A
93 |  | Error exporting Error Logs. | NO | Yes | Message POP-UP | Error exporting Error Logs. Please try again or view Error Log. | 127
94 |  | Error Log File Export done | Yes | NO | Message POP-UP | Error logs exported successfully. | 92
95 |  | Error exporting Event Logs. | NO | Yes | Message POP-UP | Error exporting Events Logs. Please try again or view Error Log. | 127
96 |  | Event Log File Export done | Yes | NO | Message POP-UP | Event logs exported successfully. | 116
110 | Media Files | Media playback file corrupted | NO | Yes | Warning POP-UP | File corrupted. Cannot play media. | N/A
111 |  | Media files deleted | Yes | NO | Message POP-UP | Selected files deleted. | N/A
112 |  | All media file deleted | Yes | NO | Message POP-UP | All media files deleted. | N/A
130 | Device Settings | RTC changed | Yes | NO | Info | N/A | N/A
131 |  | Language setting changed | Yes | NO | Info | N/A | N/A
132 |  | Maximum Video recording length updated | Yes | NO | Info | N/A | N/A
133 |  | Power saving time updated | Yes | NO | Info | N/A | N/A
134 |  | Date & Time Updated | Yes | NO | Info | N/A | N/A
150 | Battery | Battery fault | NO | Yes | Warning POP-UP | Battery fault detected. | N/A
Service required by local representative.
151 |  | Battery charger attached | Yes | NO | Message POP-UP | Charging started. | N/A
152 |  | Battery Low | NO | Yes | Warning POP-UP | Battery low. Please connect charger. | N/A
153 |  | Battery charging slower than expected | NO | Yes | Warning POP-UP | Battery charging is slow. | N/A
Check the power source.
154 |  | Battery parameters incorrect | NO | Yes | Warning POP-UP | Battery fault detected. | N/A
Service required by local representative.
155 |  | Battery charging fault | NO | Yes | Warning POP-UP | Battery charging fault. | N/A
Check the power source.
156 |  | Battery internal temperature too High | NO | Yes | Warning POP-UP | Battery temperature high. | N/A
Charging terminated.
157 |  | Battery internal temperature too Low | NO | Yes | Warning POP-UP | Battery temperature low. | N/A
Charging terminated.
158 |  | Battery < 3% device entered ship mode. | No | Yes | Warning POP-UP | Battery critically low. | N/A
Device will automatically switch-off.
159 |  | Battery secondary thermister out-of range. | No | Yes | Warning POP-UP | Internal hardware problem detected. | N/A
Service required by local representative.
160 |  | Battery charger parameters are incorrect. | No | Yes | Warning POP-UP | Internal hardware problem detected. | N/A
Service required by local representative.
161 |  | Battery Health < 70% | NO | Yes | Warning POP-UP | Battery health is reduced. Performance may be affected. Battery replacement is recommended. | N/A
Service required by local representative.
162 |  | Battery thermistor open circuit. | NO | Yes | Warning POP-UP | Battery thermistor Open Circuit. Service required by local representative | N/A
163 |  | Battery thermistor short circuit. | NO | Yes | Warning POP-UP | Battery thermistor Short Circuit. Service required by local representative | N/A
164 |  | Battery charger disconnected | Yes | NO | Message POP-UP | Charging stopped. | N/A
170 | Software upgrade | Software upgrade initiated | Yes | NO | Info | N/A | N/A
171 |  | Software upgrade succeeded | Yes | NO | Prompt POP-UP + info | Success | 68
172 |  | Software upgrade failed | NO | Yes | Prompt POP-UP + info | Error updating the software. Please try again or view Error Log. | 88
173 |  | Invalid software used for software upgrade | NO | Yes | Warning POP-UP | Software upgrade package is corrupt. | N/A
Please use correct software.
174 |  | Software upgrade is incompatible with hardware PCB. | No | Yes | Prompt POP-UP + Warning message | Software upgrade package failed. | 88
Software package not compatible with hardware version.
175 |  | Software Update taking too long | NO | Yes | Prompt POP-UP + info | Error updating the software. Please try again or view Error Log. | 88
176 |  | Software update package is older than existing | NO | Yes | Prompt POP-UP + Warning message | Software update package is older. | 88
Please use correct software.
190 | Internal Storage | Internal storage corrupted | NO | Yes | Warning POP-UP | Storage error detected. | N/A
Service required by local representative.
191 |  | Low disk space. | No | Yes | Warning POP-UP | Low disk space. Less than 30mins recording time remaining. Please delete recordings from the gallery to provide more space. | 5
200 | Device status | Device had unexpected reset | NO | Yes | Info | N/A | 75
201 |  | Internal temperature too high | NO | Yes | Warning POP-UP | Internal temperature high. | N/A
202 |  | Internal temperature too low | NO | Yes | Warning POP-UP | Internal temperature low. | N/A
203 |  | User reset the device using the power button | Yes | NO | Info | N/A | N/A
204 |  | Device auto shut-off from sleep mode | Yes | NO | Info | N/A | N/A
205 |  | Watchdog reset occurred | NO | Yes | Info | N/A | 75
206 |  | System entered sleep mode | Yes | NO | Info | N/A | N/A
207 |  | System Exit sleep mode | Yes | NO | Info | N/A | 75
208 |  | Application has exited | NO | Yes | Info | N/A | 75
209 |  | Factory reset | Yes | NO | Info | N/A | N/A
210 |  | Application checksum is invalid. | NO | Yes | Warning POP-UP | Application is corrupted. | N/A
Booting from recovery
211 |  | Software incompatible with hardware | NO | Yes | Warning POP-UP | Software incompatible with hardware. | N/A
Service required by local representative.
250 | Video Recording | Gallery button pressed while recording on | Yes | No | Message POP-UP | To view the gallery, you must first stop the recording. | N/A
251 |  | Device setting button pressed while recording on | Yes | No | Message POP-UP | To edit device settings, you must first stop the recording. | N/A
252 |  | Video Recording attempted at low battery | Yes | No | Message POP-UP | Battery charge critical. Unable to start recording. | 69
Please plug in charger.
253 |  | Battery critically low during recording | Yes | No | Message POP-UP | Battery charge critical. Recording has ended and saved. Please plug in or charge device to continue recording. | 77
254 |  | Video was split into multiple files | Yes | No | Message POP-UP | Video was split into multiple files | 90
255 |  | Recording ended & Video saved. Maximum 30 min video length reached | Yes | No | Message POP-UP | Video saved. Maximum video length reached | 94
256 |  | Recording ended & Video Saved. Maximum <x min > video length reached. Video was split into multiple files | Yes | No | Message POP-UP | Video Saved. Maximum video length reached. Video was split into multiple files | N/A
257 |  | Video recording attempt failed | No | Yes | Message POP-UP | Unable to start recording. Please retry. | N/A
258 |  | Video recording error during recording | No | Yes | Message POP-UP | Recording error occurred. File may be incomplete. | N/A
259 |  | Video recorded file corrupted / not saved | No | Yes | Message POP-UP | Recording not saved or corrupted. | N/A
300 | Touch Screen | Touchscreen sensor fault on startup | No | Yes | Warning POP-UP | Touchscreen not detected. Service required by local representative. | N/A
301 | Internal Hardware self checks | Internal isolated 3V3 Supply is faulty. | No | Yes | Warning POP-UP | Internal hardware problem detected. | N/A
Service required by local representative.
302 |  | CAM_LED6V voltage is too low. | No | Yes | Warning POP-UP | Camera LED brightness is low. | N/A
Check camera and connections.
303 |  | CAM_LED6V voltage is too high. | No | Yes | Warning POP-UP | Camera LED brightness is high. | N/A
Check camera and connections.
304 | Bronchoscope EEPROM | EEPROM not found | No | Yes | Warning POP-UP | Faulty camera connected. | N/A
Check device and connections.
305 |  | EEPROM data corrupted | No | Yes | Warning POP-UP | Faulty camera connected. | N/A
Check device and connections.
306 |  | Invalid EEPROM data | No | Yes | Warning POP-UP | Faulty camera connected. | N/A
Check device and connections.
350 | RTC | Fault reading from RTC | No | Yes | Warning POP-UP | Internal hardware problem detected. | N/A
Service required by local representative.
351 |  | Fault writing to RTC | No | Yes | Warning POP-UP | Internal hardware problem detected. | N/A
Service required by local representative.
352 | Bronchoscope ISP | Bronchoscope ISP Chip Not Detected | No | Yes | Message POP-UP | Internal hardware problem detected. | N/A
Service required by local representative.